# Lab 2: CDOM absorption						3 July 2007

# SMS 598:  Application of Remote and In-situ Ocean Optical Measurements to Ocean Biogeochemistry

## Laboratory Safety Issues – ethanol for cleaning; general laboratory safety

## Introduction

The major absorbers in seawater are water itself, chromophoric or color-absorbing dissolved organic matter (CDOM; in older literature, the term g for Gelstoff was used; the British used Gilvin), and absorbing particles. The symbol for the absorption coefficient is [a] with units of m-1, typically reported as the spectral absorption coefficient with the designation ().  In today’s lab in Station 1, we will measure A (absorbance with unitless dimensions) of CDOM with a spectrophotometer for water from three different locations; examine the blank, effect of filter size, and the potential scattering; and calculate the absorption coefficient of CDOM, aCDOM().  In Station 2, we will measure a (absorption with units of m-1) of CDOM with an ac9 with the same water; examine the pure water calibration, effect of pathlength, and the potential scattering; correct for temperature and salinity and calculate the absorption coefficient of CDOM, aCDOM().

The absorption coefficient of CDOM – aCDOM() – is operationally defined as the absorption of seawater that has been passed through a filter, typically a 0.2 m plastic filter or a G/FF filter with a nominal pore size of 0.7 m,  minus  the absorption of high quality water blank (such as produced with a Milli-Q system with a UV oxidizing cartridge).  Several things to consider are: the inclusion of viruses and bacteria in the filtered fraction, the colloidal nature of the material passing through the filter, the changing effective pore size of the size as a function of filter pad loading, the role of salts and colloidal-size particles in scattering, the quality of the blank, and the chemical nature of CDO (NB, we won’t measure dissolved organic matter, DOM, but would you expect DOM concentration to be proportional to CDOM absorbance; why?).  CDOM analyses should be carried out as soon as possible after water collection because colloid formation can continue after filtration; sampling and processing containers should be clean.

Review of Beer’s Law for a solution; both Stations 1 and 2 depend on Beer’s Law:

In   		=  I0  exp-(  c   L)

In / I0 		=  exp-(  c   L)

I0 / In		=  exp(  c   L

ln I0 / In 	=    c   L

where:

I0 is intensity of light before it passes through the sample,

In is the intensity measured at the detector after light passes through the sample,

is the molar absorption coefficient – a measure of how much light a 1 M solution of dye will absorb (m2 mole-1),

c is the concentration of the dye (mole m-3), and

L is the path length that the light must travel through the solution (m).

NB:  	In our application, the terms “  c” are combined into a single term “a”, the absorption coefficient (m-1).

Notice that the Beer’s Law equation is written in log base e (natural logarithms, ln).  However, spectroscopists historically used log base 10, rather than log base e.  The principle is the same but A, the absorbance output from the spectrophotometer, will be in log base 10.

A		=   log10 (I0/In)

=   log10e  ln(I0/In)

=   0.434 ln(I0/In)

=   0.434(a  L)

a 		= A   (0.434  L)-1    =   2.304  A  L-1

Remember from calculus, that when changing log bases: logaX = logbX logab.

To covert a natural logarithm to a base 10 logarithm, multiple by log10e (0.434).

To covert a base 10 logarithm to a natural logarithm, multiple by loge10 (2.304).

Station 1:  CDOM absorption with a spectrophotometer

Spectrophotometric measurement of filtered seawater with bench-top spectrophotometer for determining CDOM absorption coefficient and for comparison with CDOM absorption measured with ac9 or acs (Station 2, today’s lab) and CDOM fluorescence (Friday’s lab).

Goals – be able to discuss

Blanks:	does the type of water make a difference?

tap water, RO (reverse osmosis) water, MQ (Milli-Q) water

should it be filtered?

Field samples:	does CDOM absorption or its spectral slope vary?

seawater collected beyond the mouth of Damariscotta River Estuary

Damariscotta River Estuary seawater (dock)

Biscaye Pond– fresh water

Filter size:	does pore size change the result?

G/FF (nominal pore size ~ 0.7 mm) and 0.2 mm

Pathlength:	is the measurement a linear function of pathlength?  i.e., conform to Beer’s Law?

1-cm and 10-cm cells

(N.B.:  In the ac9 lab, you will use a 25-cm long pathlength).

## Scattering:	does filtered water (i.e. CDOM samples) have any detectable scattering?

## test with 1-cm cell with integrating sphere

Calculation:	convert from Absorbance (base 10) to absorption coefficient (base e)

conversion of Absorbance (unitless) determined with spectrophotometer to absorption coefficient (units of m-1) determined with ac9 or acs

# Background material on spectrophotometry

General principles of operation of spectrophotometer:

Web sites that present good reference material on the fundamentals of UV-visible spectrometry:

http://www.cem.msu.edu/~reusch/VirtualText/Spectrpy/UV-Vis/uvspec.htm#uv1

http://www.cem.msu.edu/~reusch/VirtualText/Spectrpy/InfraRed/infrared.htm

Across the Spectrum: Instrumentation for UV/Vis Spectrophotometry

Slightly modified and shortened from Shane Beck, 1998, The Scientist, 12(3): 20.

Modern spectrophotometry was pioneered by Dr. Arnold Beckman in the 1940's.

1. Light source: typical UV/Vis spectrophotometers utilize two light sources: a deuterium arc lamp for consistent intensity in the UV range (190 to 380 nm) and a tungsten-halogen lamp for consistent intensity in the visible spectrum (380 to about 800 nm).  Some spectrophotometers, such as the Cary 50, have a xenon flash lamp.

2.  Dispersion of light into different wavelengths can occur before or after the light passes through the sample.  The monochromator disperses light into different angles by prisms or holographic gratings.  NB: with a prism, the angle of dispersion can be nonlinear and sensitive to changes in temperature.  In contrast, holographic gratings eliminate nonlinear dispersion and are not temperature sensitive; they are are glass blanks with narrow ruled grooves. The grating itself is usually coated with aluminum to create a reflecting source. Grating do require filters since light is reflected in different orders with overlapping wavelengths.

Light passing through the monochromator exits as a band. The width of this band of light at half the maximum intensity is the spectral bandwidth. Bandwidth comes in to play with regard to accuracy, since the accuracy of any absorbance measurement is dependent on the ratio of the spectral bandwidth to the natural bandwidth of the substance being measured. The natural bandwidth is the width of the absorption band of the sample at half the absorption maximum. As a rule, a ratio between spectral bandwidth and natural bandwidth of 0.1 or less will generate absorbance measurements 99.5 percent accurate or better. Above this, accuracy deteriorates.

3. Light passes through the sample.  NB:  depending on the spectrophotometer, the light is either dispersed before (left) OR after passing through (right) the sample OR both, if a double monochromator system is uses. (Why might it make a difference where the monochromators are placed and where dispersion occurs.

4.  Sample absorbance is determined by comparing the intensity of the light passing through the sample and hitting the detector vs. intensity of light passing through a blank.  Detectors include: a) photomultiplier tube, with good sensitivity throughout the UV/Visible spectral range and highly sensitive at low light levels; or b) photodiode with a wider dynamic range, and consisting of a semiconductor and a capacitor to charge the semiconductor.  As light hits the semiconductor, electrons flow through it, thereby lowering the charge on the capacitor. The intensity of light of the sample is proportional to the amount of charge needed to recharge the capacitor at predetermined intervals. Often the detector is composed of a photodiode array, with photodiode detectors arranged on a silicon crystal so a spectral scan is instantaneous.

In single-beam spectrophotometers, the blank and sample are not measured simultaneously. Interspersing measurements of samples and blanks are needed to correct for lamp drift. Dual-beam spectrophotometers utilize a "chopper" or beam splitter that alternates the light path between the reference optical path and sample optical path to the detector at a speed that minimizes medium- or long-term effects of lamp drift. Some dual beam instruments scan continuously so that the sample, blank and dark reference are actually performed at different wavelengths (leading to a skewing effect as a function of wavelength, dependent upon scan speed); in others there is a phase locked wavelength drive so that the sample, blank and dark reference readings occur at the same wavelength. Blanks should be refreshed to prevent sample warming, or a cooled blank holder should be used.

If the sample is not a pure solution, scattering can occur.  An integrating sphere can be used to collect scattered light, and correct the instrument reading to provide true absorption.  The coatings on integrating spheres are highly scattering so as to ensure that the light field within the sphere is isotropic and therefore measuring a small portion of that light is equivalent to measuring it all. However, the coatings are also particularly absorptive of UV and blue radiation, which limits their utility in the UV range.

# CDOM measurement

Divide into three groups (3 students per group). Each group will perform the same measurements and scan from 200 – 800 nm:

Groups I  and II  – Cary 50 single beam, without integrating sphere,

1-cm quartz cell and 10-cm OS glass cell

Group III 		   – Cary 3E dual beam, with integrating sphere, 1-cm quartz cell

- The two Cary 50 spectrophotometers are single beam; the Cary with the integrating sphere is dual beam.

- Use fresh Milli-Q water as baseline and store; this spectrum will automatically be subtracted from all measurements.

- Don’t touch optical surface; wipe optical surface with lens paper before measurement.

- Make sure waters are at room temperature.

- Rinse cell several times with a few milliliters of sample to remove residual previous sample.

- Look through cell to ensure that there are no visual inhomogeneities (bubbles, residual mixing/turbulence between fresh and salt water sample, particles).

- Place cell in holder in exactly same orientation every time.

- Scan sample and store data.

Blanks:

1.  fresh Milli-Q

2.  RO (reverse osmosis) water

3.  tap water

Samples (2 samples each, one filtered through G/FF filter, one through 0.2 mm filter):

1.  water collected beyond mouth of DRE – G/FF and 0.2 mm filters

2. Damariscotta River Estuary water (dock) – G/FF and 0.2 mm filters

3. Biscaye Pond freshwater – G/FF and 0.2 mm filters

Assignments

1.	Does the water type affect the blank?  Which water gives the lowest blank?

2.	Subtract the Milli-Q blank and calculate the spectral absorption coefficient, aCDOM(l) for all three field samples.   Why is A dimensionless but aCDOM has units of m-1?

aCDOM(l)  =   2.304  A(l)  L-1

3. 	Does nominal filter size (G/FF ~ 0.7 mm and 0.2 mm) affect the results?

Any differences among water types?

4.   Is there any indication of scattering?  Compare absorption coefficients derived from integrating sphere vs. measurements without integrating sphere.  If there is any indication of scattering, is it more predominant at specific wavelengths?

5.   Is there a difference in absorption ~ 200 nm between fresh and saltwater samples?

If so, please characterize difference? Suggestions for why the difference?

6. 	What is the CDOM absorption coefficient at 700 and 715 nm for the three field samples? Are the values identical for all three groups and are they equal to zero?  Is there justification for subtracting the 700 or 715 nm values as a baseline?

7.  Compare a CDOM() at the ac9 wavelengths (412, 440, 488, 510, 532, 555, 650, 676).

Do the values agree?  For all samples, all wavelengths?

8.  The spectral slope of CDOM (S) can be described as

– S (l – l ref)

aCDOM()   =   aCDOM  (REF) e

where a is the amplitude of the absorption coefficient at any wavelength  (Jerlov, 1976) or at the reference wavelength, REF (usually 400, 412 or 440 nm). See Carder et al. (1989) and Blough and Del Vecchio (2002) for a discussion of the interpretation of the spectral slope.  The best method to calculate the slope is to minimize the square difference between the exponential model and the data (possibly weighed by a different error in each wavelength if the uncertainty varies as function of wavelength, e.g. due to variability in source intensity as function of wavelength). The relative (percent) error is not constant spectrally; in the red the absorption is low and the signal-to-noise high.  Slope measurements often exclude red wavelengths (e.g. the 712 nm channel in the ac9).

You may write your own code to determine the slope by non-linear exponential regression.  OR, less rigorously, you may determine the spectral slope for a CDOM() by plotting the ln-transformed values of aCDOM vs. wavelength using Excel and adding a trend line.  If you use the latter method:  Is the slope linear?  Is the slope constant over all wavelength regions.

a.  What is the spectral slope, SCDOM, for the range 412 nm – 676 nm?

b.  Does it vary among the three water types?

c.  Does is vary relative to the values obtained by with the ac9?

d.  What are the spectral slopes for other regions of the spectrum?

i.  350 – 450 nm

ii.  300 – 350 nm

# Station 2:   CDOM absorption with an ac9

In today’s lab, we will measure a (absorption with units of m-1) of CDOM with an ac9 using 0.2 mm filtered water from three different locations; examine the pure water calibration, effect of pathlength, and the potential scattering; correct for temperature and salinity and calculate the absorption coefficient of CDOM, aCDOM() (m-1); and compare with similar data measured with a bench top spectrophotometer.   Water from today’s lab will also be used to measure CDOM fluorescence in Friday’s lab (Lab 4).

Goals

Pure water:	why do we perform a pure water calibration when we have factory cals?

What type of water should be used, what other information do we need?

Field samples:	how does CDOM absorption and/or its spectral slope vary?

Seawater collected beyond the mouth of Damariscotta River Estuary

Damariscotta River Estuary water (dock)

Freshwater from Biscaye Pond

Path length:	in the spectrophotometer lab you investigate the effects of path length, how might this be simulated with an ac9 for very optically thick samples?

The ac9 you are using is 25-cm long (a 10cm long version exists as well for turbid waters) and can be thought of as the third pathlength in the series. 	A one-in-ten dilution would simulate a 2.5 cm cuvette (assuming no conformational change during dilution due to changes in ionic strength, pH etc’).

## Scattering:	does filtered water (i.e. a CDOM sample) have any detectable scattering?

## cCDOM – acCDOM = bCDOM

ac9:	How is the ac-9 cleaned? How is it stored? What type of drifts are observed with it?

Calculation:	convert measured absorption and attenuation into temperature and salinity, corrected IOPs

Apply pure water calibration, correct for temperature and salinity to arrive at the absorption and attenuation coefficients (units of m-1) for comparison with those determined spectrophotometrically

Figure 1. General schematics of WET Labs ac9.

Background material on in situ spectrophotometers such as the ac9

General principles of operation of the ac9:

The only commercially available mature in situ absorption meter is manufactured by WET Labs (Figure 1) http://www.wetlabs.com.

Some important issues related to using the ac9 can be found in Pegau et al. 1995; Pegau et al. 1997; Bricaud et al. 1995; Zaneveld et al (1994); Twardowski et al., 1999; Roesler and Boss (2007).

Schematics (Figure 2)

1. Light source: Incandescant bulb

2.  Dispersion of light into different wavelengths is done by a filter wheel with 9 filters. The filter wheel spins at 6 Hz yielding 6 spectra per second. The filter band width is 10 nm.

3. A collimated beam of light passes through the sample and onto a diffuser and a single diode detector (in the case of a; to maximize the capture of forward scattered light) or into a narrow angle detector (in the case of c, to minimize the capture of forward scattered light).

4.  Sample absorption is determined relative to a pure water calibration provided by the factory (contained in the device file, ac90nnn.dev, where nnn is the instrument serial number). Given the tendency for drift and alignment issues, it is standard practice to run your own pure water calibration prior and subsequent to each experiment.

| 1 Lamp | 6 Beam splitter |
| --- | --- |
| 2 1 mm aperture | 7 Reference detector |
| 3 6 mm aperture | 8 6 mm quartz pressure window |
| 4 38 mm singlet lens | 9 Reflective flow tube |
| 5 Interference filter | 10 Diffuser/Signal detector |

Figure 2. Schematic Representation of absorption beam optics

Figure 3. Filter wheel of acs.

The hyperspectral version of the ac9 is called the acs. Although similar in design the filter wheel holds two sections of a Linear Variable Filter (LVF), centered 180

degrees from each other on the filter wheel (Figure 3). The two filter sections are cut from a single LVF such that a portion of the spectrum around 550 nm is covered by both filters. This overlap is to allow for merging of the data from both filter sections (data generally display a gap at this overlap which needs to be corrected for). Each filter covers approximately a 72 degree section of the beam path across the filter wheel. The filter wheel rotates at a tightly controlled 8.0 rps, such that the shorter wavelength of each filter section is traversed before the longer wavelength.

# CDOM measurement

Divide into two groups (4-5 students per group). Each group will perform the exact same measurements with an instructor and ac9.

Clean the sensor prior to measurements with lens paper and ethanol. Measure the temperature and salinity of every sample.

Blanks:

1.  fresh Milli-Q

Samples:

1. Water collected beyond mouth of Damariscotta River Estuary –0.2 m filter

2. Damariscotta River Estuary water (dock) –  0.2 m filter

3. Biscaye Pond (freshwater) – 0.2 m filter

##### Assignments

1.  How do we determine Io, In, L from Beer’s Law in an ac9?

2.  Compute the median and standard deviation spectra for a and c for each set of observations. You can use provided Excel templates or write your own code.

3. Perform a pure water calibration with MQ. If your reading is significantly (absolute value of difference bigger than 0.005m-1) different from the last calibration, clean the instrument and conduct another calibration. Repeat cleaning/calibration cycle until you get at least two calibrations within 0.005m-1 of each other.

a. What is the expectation of the MQ reading relative to the device file (factory calibration)?

b. What is the reality of the MQ reading?

c. How do you use the MQ reading?

4.  Calculate the spectral absorption coefficient, a() for all field samples, applying:

a.  the pure water calibration

b.  the temperature and salinity corrections

c.  How important are the temperature and salinity corrections to your final estimates?

5.  Calculate the scattering coefficient for CDOM

a. What is your expectation of its magnitude and spectral shape?

b. What is the reality of the observations?

c.  How would you proceed with this information (i.e. perform scattering correction, what about spectrophotometric measurements)?

6.  Within the visible waveband, the spectral slope of CDOM (SCDOM) is described as:

- SCDOM ( – REF)

aCDOM()   =   aCDOM  (REF) e	       				,

where a is the amplitude of the absorption coefficient at any wavelength  (Jerlov, 1976) or at the reference wavelength, REF (usually 412 or 440 nm). See Carder et al. (1989) and Blough and Del Vecchio (2002) for a discussion of the interpretation of the spectral slope.  The best method to calculate the slope is to minimize the square difference between the exponential model and the data (possibly weighed by a different error in each wavelength if the uncertainty varies as function of wavelength, e.g. due to variability in source intensity as function of wavelength). The relative (percent) error is not constant spectrally; in the red the absorption is low and the signal-to-noise high.  Slope measurements often exclude red wavelengths (e.g. the 712 nm channel in the ac9).

You may write your own code to determine the slope by non-linear exponential regression.  OR  less rigorously, you may determine the spectral slope for a CDOM() by plotting the ln-transformed values of aCDOM vs. wavelength using Excel and adding a trend line.  (if you use the latter method, Is the slope linear?  Is the slope constant over all wavelength regions)

a. What is the spectral slope, SCDOM, for the range 412 nm – 676 nm?

b. Does it vary among the three water types?

c. Does is vary relative to the value obtained by spectrophotometry?

7.  Compare a CDOM() with those obtained with the spectrophotometer at the ac9 wavelengths. Do the values agree? Is there any spectral dependence to the differences (i.e. is there any non-random pattern to the residuals)?

###### References

Babin, M., D. Stramski, G.M. Ferrari, H. Claustre, A. Bricaud, G. Obolensky, and N. Hoepffner.  2003. Variations in the light absorption coefficients of phytoplankton, nonalgal particles, and dissolved organic matter in coastal waters around Europe.  Journal of Geophysical Research Ocean 108 (C7): article number 3211 , doi: 10.1029/2001JC000882.

Blough, N.V., and R. Del Vecchio.  2002.  Chromophoric dissolved organic matter (CDOM) in the coastal environment. In: D. Hansell and C. Carlson, Editors, Biogeochemistry of Marine Dissolved Organic Matter, Academic Press, San Diego, CA.

Bricaud, A., A. Morel, and L. Prieur. 198 1. Absorption by dissolved organic matter of the sea (yellow substance) in the UV and visible domains. Limnol. Oceanogr. 26: 43-53.

Carder, K. L, R. G. Steward, G. R. Harvey, and P. B. Ortner. 1989. Marine humic and fulvic acids: Their effects on remote sensing of ocean chlorophyll. Limnol. Oceanogr. 34: 68-8 1.

Jerlov, N.G.  1976.  Marine Optics.  Elsevier, New York.

Pegau W. S., G. Deric and J. R. V. Zaneveld. 1997. Absorption and attenuation of visible and near-infrared light in water: dependence on temperature and salinity. Applied Optics. 36: 6035-6046.

Pegau, W. S., J. S. Cleveland, W. Doss, C. D. Kennedy, R. A. Maffione, J. L. Mueller, R. Stone, C. C. Trees, A. D. Weidemann, W. H. Wells and J.R.V. Zaneveld. 1995. A comparison of methods for the measurement of the absorption coefficient in natural. J. Geophys. Res. 100(C7): 13201-13220.

Roesler, C.S., 1998: Theoretical and experimental approaches to improve the accuracy of particulate absorption coefficients derived from the quantitative filter technique. Limnology and Oceanography. 43: 1,649-1,660.

Roesler, C. S. and E. Boss. 2007. In situ measurement of the inherent optical properties (IOPs) and potential for harmful algal bloom (HAB) detection and coastal ecosystem observations, Chapter 5, pp. 153-206. In  Babin, M, C. S. Roesler and J. J. Cullen [eds.] Real-Time Coastal Observing Systems for Ecosystem Dynamics and Harmful Algal Blooms. UNESCO Series Monographs on Oceanographic Methodology.

Twardowski M. S., J. M. Sullivan, P. L. Donaghay and J. R. V. Zaneveld. 1999. Microscale quantification of the absorption by dissolved and particulate material in coastal waters with an ac-9. Journal of Atmospheric and Oceanic Technology. 16: 691-707.

Zaneveld J. R. V., J. C. Kitchen and C. M. Moore. 1994. The scattering error correction of reflecting-tube absorption meters. Proc. SPIE, Ocean Optics XII. 2258: 44-55.

Equipment and supplies

Cary 50 (2) and Cary 3E (1) spectrophotometers with external integrating sphere

1-cm and 10-cm quartz cells

ac9, power supply, comm/power cable, PC with WETView and device file

lens paper, ethanol and cleaning stick, memory sticks to save data

Water

– filtered through 0.2 m filter: seawater collected beyond mouth of Damariscotta River

Estuary, Damariscotta River Estuary water (dock), Biscaye Pond

–Blank – Milli-Qfresh but degassed, RO, tap water