Chlorophyll Fluorometer Characterization

Date:		2-04-03

Serial #:	BB2F-035

Job#:		0212021

Tech:		K.C.

CWO	(blank) 	55.6	counts

CEC        		1036	counts

# SF 	   		0.01713 ug/l per count

FSC			4121 counts

Linearity:		0.999 R2  (0–1500 counts)

0.990 R2  (0–4095 counts)

Notes:

CWO (blank) is the clean water offset value and is obtain using pure filtered de-ionized water.

CEC is the chlorophyll equivalent count.  This value is the signal output of the fluorometer when using a fluorescent proxy that has been determined to be approximately equivalent to 16.8 g/l of a Thalassiosira weissflogii phytoplankton culture.

SF is the scale factor used to derive chlorophyll concentration from the signal output of the fluorometer.  The scale factor is determine by using the following equation:  SF = (16.8) / (CEC – CWO)  e.g.(16.8 / (2865 – 238) = 0.00951)

FSC is the maximum signal count output that the fluorometer is capable of.

Chlorophyll concentration expressed in g/l (mg/m3) can be derived by using the following equation: (g/l) = (CEC – CWO) * SF

The relationship between fluorescence and chlorophyll-a concentrations in-situ is high variable.  The scale factor listed on this document was determined by using a mono-culture of phytoplankton (Thalassiosira weissflogii).  The population was assumed to be reasonably healthy and the concentration was determined by using the absorption method.  To accurately determine chlorophyll concentration using a fluorometer you must perform secondary measurements on the populations of interest. This is typically done using extraction based measurement techniques on discrete samples.  For additional information on determination of chlorophyll concentration see [ Standard Methods For The Examination Of Water And Wastewater] part 10200 H published jointly by: American Public Health Association, American Water Works Association and Water Environment Federation.