| ## Date: 11/24/09 | ## Customer: U. Of Southern Cal. | ## Repair# 7089 |
| --- | --- | --- |
| ## S/N#: ACS- 016 | ## Technician: Todd Caudill | ## Cal# 002 |

Repairs and Modifications:

Inspected the meter: the interference filters were starting to delaminate; the transmitter lenses were a little dirty; the receiver lenses were dirty; the pressure windows had minor debris on the inner surfaces.

Cleaned all optics and the interference filters.

Purged the ‘c’ receiver void.

Replaced the interference filters.

Focused and aligned the optical beams.

Stability and precision tests passed.

Shake tested the meter to 40Hz/20Hz for one hour; pressure checked for leaks.

Comments:

A complete ACS digital calibration was performed.  The new calibration sheet, the device file, and the air tracking file are enclosed.