## Slide 1: DATA PROCESSING & ANALYSIS

## Slide 2: HyperSAS and HyperPro
DATA ANALYSIS
ADDITIONAL DATA PROCESSING
Dark
Corrections
Interpolate to
bins of wavelengths
Quality Control (tilting corrections, etc)
Align data
in time
Comparison with
other radiometric
sensors
Calculate
median and std
Determine AOPs
Closure Analysis
Raw Data
Level 2
Processed Data
SOFTWARE: SatView/SatCon
Convert from counts to physical units
K
PAR
R
Rrs
Notes: After returning from the field, there are many steps to transform the raw data into a useable product. Here we show a general schematic for processing HyperSAS and HyperPro data.

## Slide 3: WISP
Raw Data
DATA ANALYSIS
DATA PROCESSING
Dark
corrections
Interpolate to bins of wavelengths
Quality Control
Align data
in time
Comparison with
other radiometric
sensors
Calculate
median and std
Determine AOPs
Closure analysis
Level 2
Processed Data
Send raw data directly to the WISP manufacturer
K
PAR
R
Rrs
Notes: The analysis for the WISP data is similar but you will need to send out the raw data to the manufacturer and there is also no dark correction.

## Slide 4
Raw data in digital counts
Instrument calibration files
Processed data in physical unit
Raw Data
Level 2
Processed Data
SOFTWARE
Notes: Using SatCon you will convert the raw data from counts to physical units.

## Slide 5
Two files are created for each sensor, one containing all dark measurements and the other, the light measurements
The dark measurement is used as baseline. Subtract these values from the light spectrum.
DATA PROCESSING
Dark
Corrections
Interpolate to bins of wavelengths
Quality Control
Align data
in time

## Slide 6
Linear interpolation to the wavelength of one of the sensors
DATA PROCESSING
Dark
Corrections
Interpolate to
bins of wavelengths
Quality Control
Align data
in time

## Slide 7
Plot data over time to identify inconsistent data that should be deleted
DATA PROCESSING
Dark
Corrections
Interpolate to bins of wavelengths
Quality Control
Align data
in time

## Slide 8
Example of the effect of the tilt angle on Rrs
DATA PROCESSING
Dark
Corrections
Interpolate to
bins of wavelengths
Quality Control
Align data
in time

## Slide 9
Remove the data with tilt angles > 5o
DATA PROCESSING
Dark
Corrections
Interpolate to
bins of wavelengths
Quality Control
Align data
in time

## Slide 10
This figure shows an example of disagreeing sensor measurements. These may be due to strong wave conditions at the time of measurement, operator error, etc.
DATA ANALYSIS
Comparison with
other radiometric
sensors
Ed
Lsky
Lsfc
This figure shows an example of more consistent and stable measurements.
It is important to determine the offset between the different instruments, if any. To do this, take measurements under controlled light conditions (for example, in a laboratory with a known light source).

## Slide 11
This example shows the median results for Lw and Rrs spectra taken over a 10 minute period and the standard deviation for the three different data sets.
DATA ANALYSIS
Calculate
median and std
Ed
Lw

## Slide 12
Remote Sensing Reflectance:
HyperSAS & WISP:
Rrs =
HyperPRO: Lee Method
Rrs=
DATA ANALYSIS
Determine AOPs
ρ= f(sky conditions, wind speed, solar zenith angle, and viewing geometry).
For ϴ =40o, Ф= 135o from the Sun and wind speed < 5m/s  ρ≈0.028 (Mobley, 1999).

## Slide 13: Differences between instruments’ response could be explained by:
Offsets between sensors,
Spatial variations during deployment,
Not accurate information about atmospheric conditions and sea level state,
Inappropriate angles corrections,
Handling errors (especially for WISP),
…
Incorrect values of sea-surface reflectance factor, ρ
DATA ANALYSIS
Closure Analysis

## Slide 14
On the left panel is shown the Rrs spectra after the offset corrections between the Ed sensors.
On the right panel, a value of ρ=0.021 was used for determining Rrs in WISP and HyperSAS sensors (corresponding to none wind speed and smaller (than 40o) zenith angles.
DATA ANALYSIS
Closure Analysis
Normalized Rrs
Normalized Rrs