## Slide 1: Radiometry Updates 7/29

## Slide 2: Dock Test 7/27
Test 1: All Sensor Test
  ‘Same’ dock location for WISP, SAS base, and buoyed HyperPro
  SAS Ed sensor located on upper dock
Test 2: Ed Sensor Test
  Same dock location for all sensors
Test 3: Water Splash Test on HyperPro
  Ran HyperPro for 1-2 minutes normal
  Ran HyperPro for 1-2 minutes after splashing water on the upwelling radiance sensor

## Slide 3: WISP Test 1: Every 30 seconds for 5-6 minutes
Good Agreement with Time
(Some boat wakes and waves)
‘Kick’ in 750-800 Smaller

## Slide 4: Test 2: Ed
Test 1: All Sensors
Good Agreement of Ed values with Time
Used Median Values To Compare with SAS/Pro

## Slide 5: Test 1: SAS Ed Sensor On Upper Dock
Difference Can Be +10%

## Slide 6: Test 2: SAS Ed Sensor On Lower Dock
Difference ½ As Big

## Slide 7: Test 3: Ed Pre-Splash Test (Some Waves)

## Slide 8: Test 1: Variability in Lw and Rrs

## Slide 9: Test 3: Variability in Lw and Rrs
Slightly Better with Ed Sensor Closer

## Slide 10: Test 1: Variability in Lt and Lsky

## Slide 11: Test 3: HyperPro Splash Test

## Slide 12: Let’s Make It More Complicated-Immersion

## Slide 13: Let’s Make It More Complicated-ρsky=0.09

## Slide 14: Recommendations?
Keeping in mind that the SAS and/or Pro may have calibration issues, using WISP in open water doesn’t look promising ‘as is’*
Gimbal the WISP
Put a level bubble on the under side of the WISP
Add automatic sampler feature (every x seconds)

## Slide 15: EXTRA FIGURES:WISP