## Slide 1: LISST calibration and measurements
Aya, Grace, Tom
Ocean Optics Summer 2013 Final Project
August 2nd, 2013

## Slide 2: Outline
Calibration
Measurements
  Lab
  Cruise
  Coccolithophore experiment
Instructional Video

## Slide 3: Calibration: Methodology
Calibration relies on agreement between measured and modeled VSF of spherical particles
Notes: Beta modeled, X_i is what we’re looking for, counts and zscat are values from machine

## Slide 4: Calibration: Methodology
Methodology from Slade and Boss 2006, “Calibrated near-forward volume scattering function obtained from the LISST particle sizer”
Measure spherical beads of known size
(for our experiment, sizes were 1μm, 20μm, 50μm)
Only use runs in which the relative error of beam c is small
Select data from particular detector rings for different size classes
  Exclude the first six rings for 1μm runs
  Exclude rings 15 to 32 for 20μm runs
  Exclude rings 11 to 32 for 50μm runs
Notes: Relative error of beam c may be high when the laser power is weak.

## Slide 5: Calibration: Coefficients
From Slade and Boss 2006
Our bead calibration
Notes: Same instrument; Should not match in magnitude because it has been factory calibrated since the paper was published

## Slide 6: Calibration: Coefficients
Our bead calibration
Calibration experiments were done with expired beads after using vortex machine on the bottle
Expiration dates:
April 2007 for 1.0 micrometer
March 2012 for 20 micrometer
July 2007 for 50 micrometer
It still works!
Notes: Same instrument; Should not match in magnitude because it has been factory calibrated since the paper was published
Read up on why there is a drop off in large rings, in paper
Error bars … maybe because of clumping?

## Slide 7: Calibration: Coefficients
Recall…
}
Output from LISST data
Obtained by bead calibration

## Slide 8: Measurements: from Lab
VSFs for lab measurements

## Slide 9
PSD for lab measurements: Volume concentration
Notes: Platy = spherical
Chaet = chains

## Slide 10: Measurements: from Lab
PSD for lab measurements: Number frequency assuming spherical shape
Notes: PSD

## Slide 11: Measurements: Cruise

## Slide 12
VSFs for cruise stations
depth = 11m
beta (m-1 sr-1)
beta (m-1 sr-1)
beta (m-1 sr-1)
beta (m-1 sr-1)
angle (radians)
angle (radians)
angle (radians)
angle (radians)
depth = 42m
depth = 32m
depth = 22m

## Slide 13: Measurements: Cruise
PSDs for cruise stations: volume concentration

## Slide 14
PSD for Damariscotta River Station: Number frequency assuming spherical shape

## Slide 15
PSD for Ocean Station : Number frequency assuming spherical shape

## Slide 16: Measurements: Coccolithophores
Investigation of changes in river coccolithophore culture scattering and size distribution pre- and post- acidification
Pre-acidification: pH 9.0
Post-acidification: pH 4.5
http://ina.tmsoc.org/CODENET/GuideImages/COpleurocs/source/40-15.htm
Notes: Acidified with acetic acid

## Slide 17: Measurements: Coccolithophores
Changes in VSF
log-log
linear
Notes: VSF; note change in forward scatter power; particles are getting smaller

## Slide 18: Measurements: Coccolithophores
Changes in PSD : volume concentration
Mean: 9μm
Mean: 8μm

## Slide 19: Measurements: Coccolithophores
Changes in PSD : number frequency assuming spherical shape
Median: 5.12 μm
Median: 6.14μm
Notes: Healthy cultures shed plates, so there’s a bimodal distribution; The acidified culture has dissolved the smaller particles.

## Slide 20: Video

## Slide 21: Acknowledgements
For just the coccolithophore experiment:
  Barney Balch, Meredith White, Dave Drapeau at Bigelow for the coccolithophore culture
  Emmanuel Boss for the LISST instrument & instruction
  Mary Jane Perry for instruction & supplies
  Nathan Briggs for guidance
  Larry Mayer for pH paper
  Ken Voss for transporting the sample