## Sheet1
2015-November-14 |  | AT32
Deck Calibration for SeaBird C-Star Transmissometer
A0 = Vair = voltage in air |  |  | 4.789
A1 = voltage w/ light path in air |  |  | 4.687
Y0 = Vd = voltage with beam blocked |  |  | 0.058
Y1 = voltage w/beam blocked deck |  |  | 0.0562
TW = 100% Transmission in pure water |  |  | 100
W0 = Vref = voltage in pure water |  |  | 4.681
M=(Tw/ [W0-Y0])*(A0-Y0)/(A1-Y1)
M= | 22.099
B= -M*Y1
B= | -1.242
CALIBRATION USING DISTILLED WATER
V(dark) = voltage w/beam blocked deck |  |  | 0.0562
V(DI) = voltage w/beam thru distilled water |  |  | 4.6288
M ={1 / [ V(DI) - V(dark)]} * 100
M= | 21.8693
B = -{ V(dark) / [ V(DI) - V(dark) ]}*100
B= | -1.2291