Protocol for processing radiometry data from Satlantic HyperPro in buoy mode (TSRB)

Emmanuel Boss and Alison Chase, UMaine, March 2026

Data was collected in buoy mode with a Satlantic HyperPro floating far from the vessel (~100m) measuring tilt and upwelled radiance (Lu) at about 17cm below the surface. A reference radiometer (Ed(0+)) was installed on the buoy.

Sensors: Ed – HyperOCR370. Lu – HyperOCR302.

Using the manufacturer software (SatCon, http://www.seabird.com/software/satcon) we converted raw files into readable text using factory calibration files dating from 2024-10-29.

We used the ‘apply immersion coefficients’ in the SatCon settings for the upwelled radiance radiometer (Lu).

We processed the resulting .dat files using MATLAB scripts

There are several steps to processing the data.

We used a 5 degrees threshold on both pitch and roll for Es and Lu data.

The basic processing steps completed by the code are:

-> Subtract time-interpolated Ed dark data from Ed light data; repeat for Lu dark and light and Eu dark and light.

-> Remove Ed, Eu and Lu spectra that have tilt values above a user-defined or default threshold (default is 5 degrees)

-> Remove Ed, and Lu spectra that are outside the middle 50th percentile of Ed data (this removes any outliers in Ed, e.g. from small clouds passing over – data are between the 25th and 75th percentliles)

-> Calculate Rrs:

Rrs (remote-sensing reflectance) is defined as the water-leaving radiance (Lwλ, i.e. the radiance just above the sea surface) normalized by the downwelling irradiance (Ed):

Rrsλ= Lwλ,0+Ed (λ,0+)

In order to compute Lwλ,0+ we need to extrapolate from the depth of the sensor (z) and cross the air-water interface:

Lwλ,0+= Luλ,zexp⁡z∙KLuλtnsw

Where t=0.98 is a transmission coefficient and nswthe index of refraction of seawater that we get from Quan and Fry (1995).

The attenuation is computed from IOPs as follows:

KLuλ= aswλ+agλ+apλ+bb,swλ+bb,pλ0.5

See chase et al., 2017 for more details.

The equation to extrapolate Lu to just below the sea surface requires information on the absorption properties of the surface water.

a_p is computed AC-s measurements of the inline system. For bb_p we used a relationship between turbidity and bbp(700) from Boss et al., 2009, and the spectrum based on Boss et al., 2013:

bb,pλ=0.0163*Turb*λ700-0.73  m-1

For CDOM we used a linear relationship established with data from Tara Europa between FDOM and CDOM at every wavelength.

The absorption by saltwater (aswλ) is also used and is taken from the recent IOCCG (2018) report augmented with temperature and salinity effects based on Sullivan et al., 2006 and Zhang et al., 2009.

We provide our measurements of Luλ,z,  Edλ,0- and their uncertainties based on the standard error of the mean. I also provide Lwλ,0+ and Rrsλ for which we propagate the uncertainties using standard propagation rules.

All raw data, calibration files, and processing codes used are in the compressed folder HyperPro_Perseverance_2026.

References:

Boss, E., L. Taylor, S. Gilbert, K. Gundersen, N. Hawley, C. Janzen, T. Johengen, H. Purcell, C. Robertson, D. W. Schar, G. J. Smith, M. N. Tamburri, 2009. Comparison of inherent optical properties as a surrogate for particulate matter concentration in coastal waters. Limnology and Oceanography: Methods 7, pp. 803-810.

Chase, A. P., E. Boss, I. Cetinic and W. Slade, 2017. Estimation of Phytoplankton Accessory Pigments from Hyperspectral Reflectance Spectra: Toward a Global Algorithm. Journal of Geophysical Research: Oceans, 122. https://doi.org/10.1002/2017JC012859

IOCCG Protocol Series (2018). Inherent Optical Property Measurements and Protocols:

Absorption Coefficient, Neeley, A. R. and Mannino, A. (eds.), IOCCG Ocean Optics and

Biogeochemistry Protocols for Satellite Ocean Colour Sensor Validation, Volume 1.0,

IOCCG, Dartmouth, NS, Canada.

Sullivan J. M., M. S. Twardowski , J. R. Zaneveld, C. Moore, A. Barnard, P. L. Donaghay, and B. Rhoades, 2006: The hyper-spectral temperature and salinity dependent absorption of pure water, salt water and heavy salt water in the visible and near-IR wavelengths (400–750 nm). Appl. Opt., 45: 5294–5309.

Quan X., and E. S. Fry (1995), Empirical equation for the index of refraction of seawater, Appl. Optics, 34, 3477–3480.