## Sheet1
SN: 1379 | Brd #188
initials | date | Process step
End Cap Assembly
MGM | 2018-12-11 | End cap clean and inspected for damage.
For PAR 1000m and PAR 7000m
MGM | 2018-12-11 | Inspect bulkhead connector
MGM | 2018-12-11 | Connector o-ring installed
Loctite applied to connector
MGM | 2018-12-11 | Connector torqued to 35 in-lb (brass) or 53 in-lb (titatnium)
PCB installation
MGM | 2018-12-11 | PCB installed with 2-56 set screw
Loctitie applied to mounting screw
MGM | 2018-12-11 | Bulkhead end cap wiring complete
MGM | 2018-12-11 | Detector installed using jig
Initial functional tests-SERIAL OPTION
Telemetry from sensor
|  | Current draw at 12V (mA)____________ |  |  |  | 18 - 22 mA
Settings updated
|  | Dark Counts________________________ |  |  |  | 327680-335872
|  | Saturation counts___________________ |  |  |  | 16777216
|  | Minimum voltage output____________ |  |  |  | 0.121 - 0.129 V
|  | Maximum voltage output____________ |  |  |  | 3.996 - 4.004 V
Temperature sensor tested
Accelerometer tested
Initial Function tests - LOG OPTION
MGM | 2018-12-11 | Current draw at 12V (mA)____________ |  |  | 18.6ma | 16 - 20 mA
MGM | 2018-12-11 | Dark voltage output_________________ |  |  | .129V | 0.121 - 0.129 V
MGM | 2018-12-11 | Saturation voltage output____________ |  |  | 4.00V | 3.996 - 4.004 V
Initial Functional tests - LIN OPTION
Telemetry from sensor
|  | Current draw at 12V (mA)____________ |  |  |  | 18 - 22 mA
Settings updated
|  | Current draw at 12V (mA)____________ |  |  |  | 16 - 20 mA
|  | Dark voltage output_________________ |  |  |  | 0.121 - 0.129 V
|  | Saturation voltage output____________ |  |  |  | 3.996 - 4004 V
FINAL ASSEMBLY
End cap potted
MGM | 2018-12-11 | Housing clean and inspected for damage
MGM | 2018-12-11 | Cosine collector installed with o-ring |  |  |  | COL-52; 17-PHI-AC
MGM | 2018-12-11 | End cap o-ring inspected and lubricated
Housing purged with nitrogen
MGM | 2018-12-11 | End cap tightened until fully seated
PAR-AUV only- loctite on end cap screws
Final functional tests - SERIAL OPTION
Telemetry from sensor
|  | Current draw at 12V (mA)____________ |  |  |  | 18 - 22 mA
|  | Dark Counts________________________ |  |  |  | 327680 - 335872
|  | Saturation counts___________________ |  |  |  | 16777216
|  | Minimum voltage output____________ |  |  |  | 0.121 - 0.129 V
|  | Maximum voltage output____________ |  |  |  | 3.996 - 4.004 V
Temperature sensor tested
Accelerometer tested
Final functional tests - LOG OPTION
MGM | 2018-12-11 | Current draw at 12V (mA)____________ |  |  | 18ma | 16 - 20 mA
MGM | 2018-12-11 | Dark voltage output_________________ |  |  | .125V | 0.121 - 0.129 V
MGM | 2018-12-11 | Saturation voltage output____________ |  |  | 3.966V | 3.996 - 4.004 V
Final functional tests - LIN OPTION
|  | Current draw at 12V (mA)____________ |  |  |  | 16 - 20 mA
|  | Dark voltage output_________________ |  |  |  | 0.121 - 0.129 V
|  | Saturation voltage output____________ |  |  |  | 3.996 - 4.004 V
Continuity, Short, and Ground fault tests
Connector tested for pin-to-metal shorts
Connector/wires tested for pin-to-pin shorts
Set Serial number
Serial number configured
Pitch and Roll offest correction
Sensor installed in tilt calibration jig
Pitch Offset
Roll Offset
Pitch offset applied
Roll offset applied
Calibration
|  | # | Date | Type (FEL/Tester) | Saturation (In-air) | Saturation (immersed) | Low-Gain Resistor
1
2
3
Pressure
|  | P (psig) | Start time |  | Weight (g) | End time |  | Weight (g)
Inspected housing for scratches
Par data and pressure logs archived
Final test and configuration
Reality check (inter-comparison) agains PAR s/n
PAR values agreement (<5% diff)
Reality check data archived