LEG7 Underway Logbook
Ali Chase, alisonpchase@gmail.com
September 23
23:22 UTC - started pump 2 and pCO2. Set in auto mode with settings of 10-10-60-10. We are at the anchorage for the site of station 43 that will begin mid-day tomorrow during the ebb tide.
Sept 24
11:31 UTC - Started pump 1 and TSG. There was a lot of excess overflow coming to the orange funnel next to the pCO2. The pumps and instruments were off and on for the next two hours while the problem was diagnosed. Also during this time the filter for the ATMOS intake to the pCO2 black box was removed by cutting the tube, water was blown out of the filter, and it was reattached.
13:21 UTC - both pumps and the pCO2 and tsg restarted. The water pulled by pump 1 had not been flowing well but we tightened the coarse filter under the floor next to U lab, and also cleaned the filter in the pump guard which was quite dirty. Everything seems to be flowing well now and the strange noise is gone from the valve tree area.
Note no microtops today; fully overcast and intermittant rain all day.
23:23 UTC - stopped and restarted pCO2 to change to a new filter on the ATMOS intake.
23:29 UTC - run syncToy
Sept 25
14:23 UTC - microtops starting with scan 0-9, sunny skies
pCO2 values still high for atmos reference even after filter change.
Sept 26
16:20 - microtops during a sunny patch; intermittant clouds.
16:30 - pump 1 stopped briefly to redirect water through optics and then restarted. force filter set on to eliminate bubbles.
16:40 - started ac-s, suvf, and hyperbb colecting data. started in total mode with auto-switching turned on. a side spectra look ok, c side goes down in blue wavelengths.
Sept 27-28
Station work during p.m. on the 27 and all day on the 28.
Sept 29
14:55 UTC - microtops, sunny skies
Sept 30
afternoon microtops, sunny
Oct 1
14:08 UTC - microtops, sunny
21:07 UTC - all UDW software stopped due to computer update. everything restarted.
Oct 2
12:20 UTC - pCO2 only just changed now to auto setting; on manual since restart evening of Oct 1. Also PAR was not restarted last night but going as of now. 30-35 kt winds and rough seas, maybe some instrument noise as a result in the pCO2 and some potential bubbles getting into UDW systems due to exposure of intake during ship movement.
after lunch microtops
Oct 3
14:10 UTC - microtops, some thin clouds
turquois water, shallow (less than 20 m)
17:34 UTC - noticed TSG was not recording data (orange not green); restarted inlinino. Checked files and some data missing from the file from the hour of 1600 on oct 3. otherwise looks fine.
19:13 UTC - pump 1 stopped to change filter in forepeak.
19:24 UTC - pump 1 restarted after filter change
Oct 4
15:25 UTC - microtops, clear and sunny
15:30 UTC - flow control software not responding; all software running slowly, closed and reopened pCO2 matlab software. Also ac-s inlinino stopped and closed and reopend, it was very slow to move the screen around. Still slow and flow control will not open. CLosing everything to restart computer.
15:43 UTC - ac-s restarted
15:54 UTC - finished inlinino restart but still the flow control will not open and everything is running very slowly
20:28 UTC - ac-s restarted, always showing a 'not responding' message on the inlinino software
20:33 UTC - ac-s failed to restart and inlinino not responding. pCO2 software not running; computer working too slowly for word processing, or synctoy
23:28 UTC - tried another full computer restart
23:39 UTC - completed restarting everything in inlinino and all running smoothly with no error messages. started pCO2 and running smoothly. did a synctoy backup of UDW computer with no problems. ***flow control software still does not open** otherwise everything working fine.
Oct 5
11:45 UTC - unplugged and replugged flow control usb in back of UDW computer and now is working well. discovered that the flow control software had been using a ton of memory. all ok now.
mid-day microtops, mostly sunny with some clouds
19:02 UTC - water trap installed on the pCO2 atmos line, also the the filter was blown clear of water during the process. pCO2 was running at the time but in sample mode for the duration.
pCO2 nitrogen tank outside checked today and reading is 185 bars. some concern of a leak; none detected but will continue to monitor and look more tomorrow in calmer seas.
Oct 6
14:16 UTC - microtops, sunny and clear
16:38 UTC - pump 1 and 2 stopped and data aquisition stopped for cleaning
17:46 UTC - pumps back on
17:59 UTC - inlinino started for ac-s, hyperbb and SUVF. some bubbles still passing through in ac-s. also lots of bubbles when switching to filtered at 18:01 UTC, stable after some seconds though.
Have been adjusting outflow valve under the floor by U-lab t increase flow, from near 3 LPM during to total to 4.5-5 LPM.
Oct 7
13:19 UTC - both pumps stopped for a few minutes
16:01 UTC - microtops, sunny and clear
21:44 UTC - computer restarted, opening again all software. At 21:52 all software completely opened. Checked the memory usage and the same problem was occuring with the flow control using a lot. Unplugged and replugged the usb again in back of the UDW computer and fine now.
Oct 8
11:30 UTC - huge trichodesmium bloom!
14:41 UTC - stopped for a deep rosette cast to respool winch. still tons of tricho at the surface, visible easily by eye
14:45 UTC - microtops, sunny clear
Oct 9
0900 UTC - stopping underway for arrival in Salvador