## DryLabLog
DATE | Jday | TIME | ACS | TSG | FRRF | Clean | Clean | Back | Clean | FW | Filter | COMMENTS (Saturday/Sunday in GREEN, Stopovers in PURPLE)
|  | UTC |  |  |  | FRRF | Fluo | Up | ACS | Flush | change
2009-10-15 |  | 14:00 | OFF | OFF | OFF |  |  |  |  | Y
|  | 15:00 | [9 empty] | ARRIVE BIZERTE
2009-10-16
2009-10-17
2009-10-18 |  | 11:00 | [6 empty] | Y
2009-10-19 | 291 | 8:00 | [9 empty] | DEPART BIZERTE
|  | 10:00 | ON | ON | ON
2009-10-20 | 292 | 8:00 |  |  |  | Y | [5 empty] | Forgot to put slide to 2 (FRRF) until 11:30
2009-10-21 | 293 | 7:00 | OFF | OFF | OFF |  |  |  |  | Y |  | ARRIVE NAPLES
2009-10-22
2009-10-23
2009-10-24 | 294 | 3:50 |  |  |  | Y
|  | 6:45 | [5 empty] | Y
|  | 17:15 | [6 empty] | Y
2009-10-25 | 297 | 2:45 | [9 empty] | DEPART NAPLES
|  | 3:10 | ON | ON | ON
|  | 14:25 | [9 empty] | NMEA on TSG NOT CHANGING. Tried restarting, but still not working
We are on Station, no time to fix
|  | 17:36 | [9 empty] | TSG NMEA good again
2009-10-26 | 298 | 7:45 | [9 empty] | TSG NMEA stuck again from 01:30. Restarted
TSG NMEA stuck again from 11:00 - 17:00. Restarted
|  | 13:00 | [9 empty] | Found power off in Dry Lab: new batteries installed and on since 23:00
|  | 16:24 | [9 empty] | TSG NMEA stuck again. No luck restarting.
|  | 17:14 | OFF | OFF | OFF |  |  |  |  | Y |  | ARRIVE at Anchorage
2009-10-27 | 299 | 5:20 | ON | ON | ON | [6 empty] | DEPART Anchorage
|  | 7:39 | [9 empty] | TSG NMEA stuck again. On station, no time to fix
Tried several times to restart thru the day, no luck
2009-10-28 | 300 | 3:15 | [9 empty] | Started again
|  | 7:15 | OFF | OFF | OFF | [6 empty] | No water to FRRF. Lots of air in system. Stopped Pump
|  | 10:22 | ON | ON | ON | [6 empty] | Start pump. No water to FRRF. On station, no time to fix
|  | 12:00 | [9 empty] | NMEA stopped
|  | 17:44 | OFF | OFF | OFF |  |  |  |  | Y |  | ARRIVE at Anchorage
2009-10-29 | 301 | 9:30 | [9 empty] | ARRIVE LA VALETTA
|  | 11:30 |  |  |  | Y |  |  | Y | Y | Y | New 0.2 um Filter
|  | 14:20 | [9 empty] | NMEA test: direct connection to deck unit. No errors over 8 hours
2009-10-30
2009-10-31 | 303 | [10 empty] | NMEA test: connect to splitter. ~ 7 errors over 10 hours (additional chars)
2009-11-01
2009-11-02
2009-11-03
2009-11-04 | 307 | 6:00 | [9 empty] | DEPART LA VALETTA
|  | 6:20 | ON | ON | ON | [6 empty] | Total Flow = 4.600 l/m Filtered Flow = 4.000 l/m
|  | 8:00 | [9 empty] | Sorted flow to FRRF. OK now.
2009-11-05 | 308 | 7:10 | [9 empty] | Sent TSG day 307
|  | 11:45 | OFF | OFF | OFF | [6 empty] | ARRIVE TRIPOLI
|  | 12:50 |  |  |  | Y |  |  |  | Y |  | Sent TSG day 308
|  | 14:00 | [6 empty] | Y
2009-11-06
2009-11-07
2009-11-08 | 311 | 8:00 | [5 empty] | Y |  |  |  | onto USB EGO's
|  | 8:35 | [9 empty] | Installed new Valve Control beta 3. Now remembers previous settings
2009-11-09
2009-11-10
2009-11-11 | 314 | 7:00 | [9 empty] | DEPART TRIPOLI
|  | 7:40 | ON | ON | ON | [6 empty] | NO NMEA, tried all day to fix
2009-11-12 | 315 | 5:43 |  |  |  | Y | [5 empty] | ON but NO NMEA
2009-11-13 | 316 | 8:00 | OFF | OFF | OFF |  |  |  |  | Y |  | ARRIVE LA VALETTA
|  | 3:30 | [9 empty] | DEPART LA VALETTA
2009-11-14 | 317 | 4:45 | ON | ON | ON | [6 empty] | ACS - no continuous, schedule only TSG using Tara_No_NMEA.con
|  | 5:00 | [9 empty] | NMEA data logged onto Library Dell Laptop
|  | 5:20 |  |  |  | Y | [5 empty] | Blockage in FRRF, now working
|  | 8:45 | [9 empty] | TSG back on, took serial cable for testing
2009-11-15 | 318 | 11:00 |  |  |  | Y | [5 empty] | NB GPS DATA is being saved in FRRF *.res files!!!!
2009-11-16 | 319 | 5:50 | [9 empty] | Power outage. Restarted immediatley
|  | 11:20 | [9 empty] | Power outage. Restarted immediatley
|  | 12:11 | [9 empty] | Found FRRF failed. Outage?
2009-11-17 | 320 | 18:45 |  |  |  | Y
2009-11-18 | 321 | 11:00 | [9 empty] | Power Outage, reset at 11:40
|  | 21:50 |  |  |  | Y
2009-11-19 | 322 | 8:10 | OFF | OFF | OFF | Y |  | Y | Y | Y |  | Switched out ACS but connectors different.
|  | 9:00 | [9 empty] | ARRIVE DUBROVNIK
|  | 18:00 | [9 empty] | NMEA works! Wired direct to PC com, bypass switch, using yellow + blue
2009-11-20
2009-11-21
2009-11-22 | 325 | 7:30 | [9 empty] | DEPART DUBROVNIK
|  | 8:40 | OFF | ON | ON | [6 empty] | Flow low: filt = 2.7 total = 4.8, keeping valve on
|  | 15:55 | [9 empty] | GPS alarm won't turn OFF.
|  | 17:00 | [9 empty] | Now using UPPER GPS
2009-11-23 | 326 | 6:45 | [9 empty] | TSG file 08:40 - 20:04 before stopping
|  | 6:50 |  |  |  | Y
|  | 17:03 | [9 empty] | No NMEA change in data since 09:00. Saving without NMEA.
2009-11-24 | 327 | 6:40 | [9 empty] | TSG data file only 03:00 to 09:00
|  | 18:15 |  |  |  | Y
2009-11-25 | 328 | 8:30 | [9 empty] | Found all off - Power outage; Restarted
|  | 15:00 | OFF | OFF | OFF | Y |  | Y |  | Y |  | AT ANCHOR
2009-11-26 | 329
2009-11-27 | 330 | 9:02 |  |  |  | Y
|  | 10:00 | [6 empty] | Y | Y | Y | New 0.2 um Filter
2009-11-28 | 331 | 9:00 | [9 empty] | DEPART ATHENES
|  | 10:00 | ON | ON | ON | Y
|  | 13:40 |  |  | ON | [6 empty] | FRRF water stopped moving in tube. Stopped acquisition, cleanned tubes and restart acquisition
2009-11-29 | 332 | 7:30 |  |  |  | Y | [5 empty] | restart TSG to change vizualisation settings
2009-11-30 | 333 | 7:00 | [9 empty] | No data from ACS since 16-Nov,restated acqiusition software
01-dec-09 | 334 | 8:00 |  |  |  | Y
01-dec-09 |  | 8:40 | [9 empty] | Water is not flowing in FRRF constantly. Bubbles in CTD as well
01-dec-09 |  | 18:40 | OFF | OFF | OFF | [6 empty] | ARRIVE LIMASSOL
02-dec-09 | 335 | 22:00 | ON | ON | ON | Y | [5 empty] | DEPART LIMASSOL
03-dec-09 | 336 | 13:24 |  |  |  | Y | [5 empty] | FRRF motor going in wrong direction (since yesterday?)
04-dec-09 | 337 |  |  |  |  | Y
05-dec-09 | 338 |  |  |  |  | Y
06-dec-09 | 339 | 8:00 | OFF | OFF | OFF | Y |  |  | Y | Y |  | ARRIVE LIMASSOL
[6 empty] | Y |  |  | N | ?
[6 empty] | Y |  |  | N | ?
[6 empty] | Y |  |  | N | ? |  | No report from scientists onboard
[6 empty] | Y |  |  | N | ?
[6 empty] | Y |  |  | N | ?
14-dec-09 |  | 10:35 | ON | ON | ON | Y |  |  | N | ? |  | DEPART BEYRUTH
17-dec-09 |  | 10:00 | OFF | OFF | OFF | Y |  |  | Y | Y |  | ARRIVEE PORT SAID
Suez Transit, All OFF
7th Jan | [7 empty] | Y
8th Jan |  | 5:35 | [9 empty] | Tested after layoff. All good EXCEPT ACS Laptop giving FAN ERROR
|  | 14:30 | [9 empty] | DEPART SHARM EL SHEIK
|  | 15:50 | ON | ON | ON | [6 empty] | All on and running
9th Jan |  | 4:50 | [9 empty] | All good
|  | 6:10 |  |  |  | Y
|  | 17:58 |  |  |  | Y | [5 empty] | Found slide on 3... UG
10th Jan |  | 7:00 |  |  |  | Y
|  | 17:00 |  |  |  | Y
2010-01-11 |  | 4:50 |  |  |  | Y | [5 empty] | All good!
|  | 22:19 | [9 empty] | Found TSG bombed out, restarted file as 2010_01_11_a
2010-01-12 |  | 4:30 | OFF | OFF | OFF | [6 empty] | Power OFF to Dry Lab
|  | 8:36 | ON | ON | ON | Y | [5 empty] | Power ON. ACS filtered flow = 2.7, Total = 4.8
2010-01-13 |  | 6:10 |  |  |  | Y
|  | 18:40 |  |  |  | Y
|  | 19:20 | [9 empty] | SeaFlow ON
2010-01-14 |  | 6:20 | OFF | OFF | OFF |  |  |  |  | Y
|  | 7:30 | [9 empty] | ARRIVE JEDDAH
|  | 9:30 | [5 empty] | Y | Y
2010-01-15 |  | 8:00 |  |  |  | Y |  |  |  | Y | Y | New 0.2 um Filter
2010-01-16
2010-01-17
2010-01-18 |  | 7:20 | [9 empty] | DEPART JEDDAH
|  | 7:40 | ON | ON | ON | Y | [5 empty] | All on All OK! SeaFlow ON, waiting for internet
|  | ~? |  |  |  | Y
2010-01-19 |  | 4:00 | [9 empty] | All OK
|  | 4:06 |  |  |  | Y | [5 empty] | Found TSG timed out + FRRF errors. 2 Power outages in 5 minutes
|  | 4:30 | [9 empty] | TSG timed out again. Files *_19_b.hex
|  | 16:00 |  |  |  | Y
2010-01-20 |  | 3:15 |  |  |  | Y | [5 empty] | All good
|  | 3:50 | [9 empty] | Found power off, restarted TSG + FRRF
2010-01-21 |  | 4:50 |  |  |  | Y | [5 empty] | Found power off, restarted TSG + FRRF
|  | 8:23 | [9 empty] | Found power off, restarted TSG + FRRF
|  | 16:55 |  |  |  | Y
2010-01-22 |  | 3:30 | [9 empty] | Found power off, restarted TSG + FRRF
|  | 3:50 |  |  |  | Y | [5 empty] | All restarted
|  | 15:57 |  |  |  | Y | [5 empty] | Found FRRF not reponding. Cleaned + restarted
2010-01-23 |  | 4:05 |  |  |  | Y | [5 empty] | TSG file restarted. No power cut!
|  | 16:38 |  |  |  | Y
2010-01-24 |  | 10:55 | OFF | OFF | OFF | [6 empty] | FRRF unresponsive.
|  | 11:00 | [9 empty] | ARRIVE DJIBOUTI
|  | 11:30 |  |  |  | Y |  | Y | Y | Y |  | ACS cleaned and packed for return for calibration.
2010-03-02 | [11 empty] | Depart Abu Dhabi
Replace ACS filter
2010-03-02 |  | 6:00 | ON | ON | ON | Y | [5 empty] | Many power drops => many TSG files
2010-03-03 |  | 0:00 |  |  |  | Y | [5 empty] | Many power drops => many TSG files
2010-03-04 |  | 0:00 |  |  |  | Y | [5 empty] | Many power drops => many TSG files Installed backup drive and syncback.exe
2010-03-05 |  | 4:20 | OFF | OFF | OFF | Y |  | Y | Y | Y |  | Found TSG software in default
2010-03-05 | [11 empty] | Arrivée Masqat
Vane 19 added on ACS intake
2010-03-09 |  | 15:40 | ON | ON | ON | Y |  | Y |  |  |  | Depart Masqat
2010-03-10 | [5 empty] | Y |  | Y |  |  |  | Many TSG restart
2010-03-11 | [5 empty] | Y |  | Y |  |  |  | Many TSG restart
2010-03-12 | [5 empty] | Y |  | Y |  |  |  | Many TSG restart
2010-03-13 | [5 empty] | Y |  | Y |  |  |  | Many TSG restart
2010-03-14 | [5 empty] | Y |  | Y | Y |  |  | Replace ACS filter
2010-03-15 | [5 empty] | Y |  | Y
2010-03-16 | [5 empty] | Y |  | Y
2010-03-17 | [5 empty] | Y |  | Y |  |  |  | Clean TSG, rince ACS with fresh water
2010-03-18 | [5 empty] | Y |  | Y
2010-03-19 | [5 empty] | Y |  | Y
2010-03-20 | [5 empty] | Y |  | Y |  |  |  | Flow not started on ACS PC
2010-03-21 | [5 empty] | Y |  | Y
2010-03-22 | [5 empty] | Y |  | Y
2010-03-23 |  | 5:00 |  |  |  | Y |  | Y |  |  |  | FRRF found on position 3
|  |  | OFF | OFF | OFF | Y |  | Y | Y | Y |  | Deep cleaning of TSG
Arrivée Mumbay
2010-03-28 |  | 2:00 | ON | ON | ON | Y |  | Y |  |  |  | Depart Mumbay
2010-03-29 | [5 empty] | Y |  | Y
2010-03-30 | [5 empty] | Y |  | Y |  |  |  | FRRF air intake problem solved !
2010-03-31 | [5 empty] | Y |  | Y |  |  |  | TSG acquisition using SEATERM (capture)
2010-04-01 | [5 empty] | Y |  | Y
2010-04-02 |  | 6:00 |  |  |  | Y |  | Y |  |  |  | FRRF found on position 3, most TSG data lost
2010-04-03 | [5 empty] | Y |  | Y
2010-04-04 | [5 empty] | Y |  | Y
2010-04-05 |  |  | OFF | OFF | OFF | Y |  | Y | Y | Y |  | Arrivée Male
2010-04-10 | 99 | 4:00 | [9 empty] | Depart Male
|  | 4:50 | ON | ON | ON | [6 empty] | All on + OK
|  | 5:00 | [9 empty] | FRRF OK
|  | 6:45 | [9 empty] | Power Outage
|  | 7:00 | [9 empty] | Power ON
|  | 15:30 |  |  |  | Y |  | Y |  |  |  | Found TSG NMEA stuck
2010-04-11 | 100 | 1:13 | [9 empty] | All OK!
|  | 6:08 | [9 empty] | Found FRRF failed, restarted
|  | 14:19 | [9 empty] | TSG - no NMEA, recycled interface box
|  | 19:00 |  |  |  | Y |  | Y |  |  |  | Found FFRF slide on "3" / ACS frozen screen
2010-04-12 | 101 | 3:47 |  |  |  | Y
|  | 3:53 | [9 empty] | TSG new file tsg_2010_04_12_a
|  | 13:00 | [5 empty] | Y |  |  |  | TSG - found time stamp stuck on 04:12:31
|  | 13:15 |  |  |  | Y
2010-04-13 | 102 | 1:42 |  |  |  | Y | [5 empty] | ACS - device time out error, last file at 01:39 so OK
|  | 4:43 | [9 empty] | Testing TSG / GPS
|  | 13:43 | [9 empty] | ACS - device time out error
|  | 13:53 |  |  |  | Y | [5 empty] | TSG - start data saving in Seasave file
|  | 14:04 |  |  |  | Y |  | Y |  |  |  | ACS - 3x time out errors in 10 minutes!
|  | 18:34 | [9 empty] | Found FRRF flow in wrong direction
2010-04-14 | 103 | 2:51 | [9 empty] | ACS - timed out / TSG - NO conductivity data
|  | 3:56 | [9 empty] | ACS timed out
|  | 4:43 | [9 empty] | ACS timed out. Restarted laptop also
|  | 6:47 | OFF | OFF | OFF | [6 empty] | AT ANCHOR GAN ISLAND. Work on generators needed
2010-04-15 | 104 | 15:02 | ON | ON | ON | [6 empty] | UP ANCHOR.
|  | 15:56 |  |  |  | Y | [5 empty] | *Flow Low* Flitered: 2.2 l/m Total: 2.7
2010-04-16 | 105 | 0:34 | [9 empty] | ACS timed out 6 hrs lost, on top of 12 already lost. FLOW STILL LOW
|  | 1:47 | [9 empty] | ACS timed out
|  | 1:58 | OFF | OFF | OFF |  |  |  |  | Y |  | Low Flow not good for pump + generator problems = No Underway
|  | 15:05 |  |  | ON | [6 empty] | Still trying to find Obstruction in flow. FRRF restarted
2010-04-17 | 106 | 4:02 |  |  | OFF | [6 empty] | Working on Flow Problem
|  | 6:03 | ON | ON | ON | [6 empty] | FOUND FLOW OBSTRUCTION: shells in hose just by SBE45, slight discontinuity.
Might be due to dislodged filter basket in mud box???
All instruments ON. Filtered = 4.6 l/m, TOTAL = 5.2 l/m
|  | 10:54 | [9 empty] | ACS Timed out
|  | 12:03 | [9 empty] | ACS Timed out
|  | 14:40 | [9 empty] | ACS Timed out
|  | 15:34 | [9 empty] | ACS Timed out
|  | 15:35 |  |  |  | Y |  | Y |  |  |  | Single spike in attenuation ~550um
|  | 17:23 | [9 empty] | FRRF turned to "2"
|  | 17:43 | [9 empty] | Took out ACS USB hub, Valve control + Compass now on COM5 + COM1
2010-04-18 | 107 | 0:50 | [9 empty] | ACS Timed out
|  | 4:22 | [9 empty] | ACS Timed out
|  | 14:41 |  |  |  | Y |  | Y
|  | 17:02 | [9 empty] | ACS Timed out
|  | 20:28 | [9 empty] | POWER OUTAGE: Change over Gen set, restart FRRF + TSG
2010-04-19 | 108 | 0:15 | [9 empty] | POWER OUTAGE: Change over Gen set, restart FRRF + TSG
|  | 3:20 | [9 empty] | POWER OUTAGE: Change over Gen set, restart FRRF + TSG
|  | 3:43 | [9 empty] | Low Disk Space: backed up + checked on backup drive PLUS onto EGO before deleting
|  | 4:08 | [9 empty] | ACS timed out
|  | 7:45 | [9 empty] | POWER OUTAGE: Change over Gen set, restart FRRF
|  | 15:25 |  |  |  | Y |  | Y
2010-04-20 | 109 | 4:33 |  |  |  | Y | [5 empty] | All OK!
|  | 12:04 | [9 empty] | TSG GPS testing back on at 12:13
|  | 15:35 |  |  |  | Y |  | Y
2010-04-21 | 110 | 7:21 | [9 empty] | TSG GPS testing
|  | 7:37 | [9 empty] | ACS went wild, restarted laptop
|  | 8:05 | [9 empty] | backup started to TaraNAS, TSG + then ACS off
|  | 11:08 | [9 empty] | ACS back on
|  | 16:14 |  |  |  | Y |  | Y
2010-04-22 | 111 | 4:55 | [9 empty] | ACS timed out
|  | 6:19 |  |  |  | Y
|  | 7:56 | [9 empty] | Found ACS screen blank, dont know why, but programs still running
|  | 10:18 | [9 empty] | ACS timed out
|  | 11:02 | [9 empty] | Found ACS screen blank, dont know why, but programs still running
|  | 13:05 | [9 empty] | ACS timed out
|  | 13:54 | [9 empty] | ACS timed out
|  | 15:40 |  |  |  | Y |  | Y
|  | 16:32 | [9 empty] | ACS timed out
2010-04-23 | 112 | 4:32 | [9 empty] | ACS timed out
|  | 5:30 | OFF | OFF | OFF | Y |  | Y |  | Y |  | Noticed GPS wire to FRRF fallen out. No GPS since day 106. Bugger.
ARRIVE Anchorage St. Brandon
2010-05-08 | 126 | [10 empty] | DEPART MAURITIUS
|  | 6:30 | ON | ON | ON | [6 empty] | Lots of bubbles on start up. All FRRF tubing very Black!
|  | 6:48 |  |  |  | Y | [5 empty] | replaced tubing running into cuvette of FRRF
|  | 7:43 | [9 empty] | ACS screen OFF, continuous problem all day. Will not come on without restarting laptop.
|  | 15:12 | [9 empty] | Generator change over, restart FRRF
|  | 17:32 |  |  |  | Y |  | Y
ACS screen savers on "never", but keeps blacking out. Heat?? Continues to work though if left
Found a way to restart: Fn+F4 (sleep), then power on will restart screen
2010-05-09 | 127 | 2:47 | [9 empty] | All OK
|  | 15:54 |  |  |  | Y |  | Y
|  | 16:14 | [9 empty] | Generator change over, restart FRRF
2010-05-10 | 128 | 6:17 | [9 empty] | ACS timed out, not much lost. Spike in attenuation at 560nm on restart
|  | 8:54 | [9 empty] | Found FRRF bombed out
|  | 13:54 |  |  |  | Y |  | Y
|  | 14:17 | [9 empty] | Power outage, restarted FRRF
2010-05-11 | 129 | 2:08 | [9 empty] | All OK
|  | 14:41 | [9 empty] | Generator change over, restart FRRF
2010-05-12 | 130 | 5:39 | OFF | OFF | OFF | Y |  | Y | Y | Y
ARRIVE Reunion
2010-05-15 | 133 | 13:30 | [9 empty] | DEPART Reunion
|  | 14:18 | ON | ON | ON | Y | [5 empty] | Sasha Tozzi here, replaced FRRF, new tubing….
|  | 15:30 | [9 empty] | Power outage
|  | 15:42 | [9 empty] | Restart all. ACS big spike attenuation ~570nm
|  | 16:36 | [9 empty] | Found ACS Timed Out, restarted with no spike
|  | 18:23 | [9 empty] | Genrator changed, FRRF restart
2010-05-16 | 134 | 14:25 |  |  |  | Y |  | Y
|  | 14:47 | [9 empty] | Power outage, restart TSG + FRRF
|  | 15:19 | [9 empty] | Found FRRF failed, restarted
2010-05-17 | 135 | 1:40 |  |  |  | Y | [5 empty] | All good. Couldn't wake ACS had to restart laptop
|  | 2:45 | [9 empty] | Power outage, all restarted
|  | 2:55 | [9 empty] | Power outage, ACS restarted (with Spike at ~570nm)
|  | 19:42 | [9 empty] | Power outage, all restarted. ACS not logging since 17:56
|  | 19:52 | [9 empty] | Power outage, all restarted. ACS no spike
2010-05-18 | 136 | 3:36 |  |  |  | Y | [5 empty] | All OK
|  | 6:55 | [5 empty] | Y |  |  |  | Founs ACS laptop off. Restarted ACS
|  | 14:46 | [9 empty] | Generator change, FRRF restart
|  | 14:59 | OFF | [8 empty] | Collars not seated well on ACS… has been leaking into bilge. Turned off ACS
|  | 15:11 | ON | [8 empty] | ACS on
2010-05-19 | 137 | 2:10 |  |  |  | Y | [5 empty] | All OK exccept FRRF zero point error
|  | 8:19 | [9 empty] | Power outage ~20 mins ago
|  | 8:54 | [5 empty] | Y |  |  |  | Power outage. TSG restart
|  | 14:26 | [9 empty] | FRRF float pt 0 zero. ACS timed out.
2010-05-20 | 138 | 3:39 |  |  |  | Y | [5 empty] | All OK
|  | 5:23 | [5 empty] | Y
|  | 7:37 | [9 empty] | Power outage, restart
|  | 8:17 | OFF | OFF | OFF |  |  | Y |  | Y |  | ARRIVE Diego Suarez
2010-05-23 | 141 | 11:35 | ON | ON | ON | Y |  |  |  | Y |  | DEPART Diego Suarez
2010-05-24 | 142 | 3:00 |  |  |  | Y | [5 empty] | All OK
|  | 15:26 | [9 empty] | B/U TSG + ACS
|  | 16:22 | [5 empty] | Y |  |  |  | Generator change restart TSG + FRRF
2010-05-25 | 143 | 3:59 |  |  |  | Y | [5 empty] | ACS restarted
|  | 10:01 | [9 empty] | Total=4.4 Filtered=3.8, PUMP STOPPED on station while trying to clear blockage(?) Was HOT
Installed "spare pump" but is loud + very high flow (Total=5.7, Filtered=4.8). Installed valve to restrict flow…
|  | 13:36 | [9 empty] | Spare pump soon got hot HOT. (Turned out hot normal here in hot waters!)
|  | 16:32 | [9 empty] | FINALLY GOT PUMP WORKING AGAIN. Julien Engineer opened up old one + its working now… maybe cleared a blockage?
|  | 18:00 | [5 empty] | Y |  |  |  | Power outage + restart
2010-05-26 | 144 | 3:30 |  |  |  | Y | [5 empty] | All OK! Flow back to normal!
|  | 8:04 | OFF | OFF | OFF |  |  | Y |  |  |  | ARRIVE ANCHORAGE Mayotte
23-june_10 | [5 empty] | Y |  |  | Y | Y |  | TSG cleaned
24_june_10 | 174 | 12:00 | [9 empty] | DEPARTURE MAYOTTE
24_june_10 | 174 | 12:20 | ON | [7 empty] | Y | new filter on ACS, total =5 filtered =4.8; new laptop
|  | 12:40 |  | ON | [7 empty] | TSG with new GPS , yyyy corrupted in NMEA , but seasave OK on 1998;interface on 24VDC
|  | 13:00 |  |  | ON | [6 empty] | filter #2 ( value declared in protocol is : 1, wrong)
|  | 14:00 | [9 empty] | TSG :sesave buggs , many trials then capture with Seaterm
25-june-10 | 175 | 5:28 | [9 empty] | FRRF : filter #3 ( value declared in protocol is : 1, wrong), TSG ok on seaterm
|  | 11:00 |  |  |  | Y |  | Y |  |  |  | FRRF : filter #1 ( value declared in protocol is : 1, correct)
26-june-10 | 176 | 8:00 | [9 empty] | sent 5mn reduced for JJ175 (incomplete :00 to 17 UTC)
|  | 14:35 |  |  |  | Y |  | Y |  |  |  | Syncback installed onACS laptop
27-june-10 | 177 | 8:00 | [9 empty] | sent 5mn reduced for JJ176 (incomplete :05 to 24 UTC)
28-june-10 | 178 | 7:00 |  |  |  | Y |  | Y |  |  |  | Juan de Nova mooring, sent 5mn reduced for JJ177
29-june-10 | 179 | 3:44 | OFF | OFF | OFF | [6 empty] | stopped all,sea water pump failure (brushes), put spare pump, flow to zeros since midnight????
|  | 4:55 | ON | ON | ON | [6 empty] | restarted all
|  | 8:00 |  |  |  | Y |  | Y
30-june-10 | 180 | 8:00 |  |  |  | Y |  | Y |  |  |  | sent 5mn reduced for JJ178, not 179 ( corrupted)
2010-07-01 | 181 | 8:00 |  |  |  | Y |  | Y |  |  |  | sent 5mn reduced for JJ179, and 180
2010-07-02 | 182 | 21:41 |  |  |  | Y |  | Y |  |  |  | restarted FRRF
2010-07-03 | 183 | 4:52 | [9 empty] | restarted FRRF
|  | 7:00 | OFF | OFF | off /on | Y |  | Y | Y | Y |  | stopped ACS TSG FRRF for cleaning, changed acs FILTER ( moored at Europa)
2010-07-04 | 184 | 11:00 | ON | ON | ON | [6 empty] | started TSG ACS for CTD cast at mooring , then out to sea
5_jul-10 | 185 | 16:15 |  |  |  | Y |  | Y |  |  |  | restarted FRRF,sent TSG 181/182
2010-07-06 | 186 | 3:54 | [9 empty] | restarted ACS
|  | 6:00 | [9 empty] | power outage , restarted FRRF
2010-07-07 | 187 | 7:00 |  |  |  | Y |  | Y
2010-07-08 | 188 | 5:00 |  | OFF | OFF | [6 empty] | TSG (spare) pump failure during the night
|  | 6:00 |  | ON | ON | [6 empty] | old (repaired ) pump back in place :OK
|  | 8:50 |  |  |  | Y |  | Y
|  | 15:50 | [9 empty] | Frrf Gps failure , reboot
2010-07-09 | 189 | 6:20 | [5 empty] | Y
2010-07-10 | 190 | 7:20 | [9 empty] | reboot ACS ( vane control problem),reboot FRRF (power outage)
|  | 08::00 |  |  |  | Y |  | Y
2010-07-11 | 191 | 8:37 | [9 empty] | reboot ACS ( display locked)
|  | 11:00 |  |  |  | Y |  | Y
|  | 12:44 | [9 empty] | reboot ACS ( display locked)
2010-07-12 | 192 | 8:20 | [5 empty] | Y | [5 empty] | s
|  | 13:40 |  |  |  | Y
2010-07-13 | 193 | 00 ??? | OFF | OFF | [7 empty] | 24VDC power outage ( fuse 13/14 in EE01)
|  | 5:00 | ON | ON | [7 empty] | restarted ACS+TSG, new TSG file (100713)
|  | 8:00 |  |  |  | Y |  | Y
2010-07-14 | 194 | 9:30 | [9 empty] | ACS frozen, restarted
|  | 10:00 |  |  |  | Y |  | Y
2010-07-15 | 195 | 7:30 |  |  |  | Y |  | Y
2010-07-16 | 196 | 10:15 | OFF | OFF | OFF | Y |  | Y | Y | Y |  | ARRIVAL CAPE TOWN
SIMON'S TOWN DRY DOCK - Dismantle underway system
2010-08-27 | 239 | [10 empty] | SIMON'S TOWN DRY DOCK - Install underway system
2010-08-28 | 240 | [9 empty] | Y | Reinstalled Underway. New 0.2 um Filter. SBE45 not seen by Interface Box | Interface Box TSG problem. Turned out to be baud rate mismatch
2010-08-30 | 241 | [10 empty] | Found SBE45 baud rate = 9600 and | ALSO discovered Toughbook serial port no good for sending commands
[12 empty] | Tested GPS O/P on QCS laptop overnight, dropped out after several hours | Interface Box GPS problem, turned out to be dodgy connector
DEPARTURE SIMON'S TOWN
2010-09-01 | 242 |  | ON | ON | ON | [6 empty] | Sea Trials: Simons Town to Cape Town. All running OK | Connector re-soldered, but problem persists, internal connection bad?
|  | 10:52 | [9 empty] | ACS losing records but NO ERROR MSG. Restarted
|  | 16:06 | OFF | OFF | OFF | [6 empty] | END + FW flush | Marc working on independent means to collect GPS data ditrectly
ARRIVAL CAPE TOWN
DEPARTURE CAPE TOWN
2010-09-05 | 247 | 9:15 | ON | ON | ON | [6 empty] | ACS + FRRF OK, TSG on motocross. Flow 4.9 / 4.3
|  | 12:44 | [9 empty] | ACS losing data, restarted compass
|  | 13:09 | [9 empty] | ACS losing data, restarted compass
|  | 17:08 | [9 empty] | ACS losing data, restarted compass
|  | 17:43 | [9 empty] | FRRF found on "0" reset to "3"
|  | 19:25 |  |  |  | Y |  | Y
|  | 19:45 | [9 empty] | Found FLOW LOW. Seacock valve incorrectly set. Press variations must've cleared blockage
Seacock valve mis-labelled during refit.
2010-09-06 | 248 | 2:41 | [9 empty] | ACS "device timed out" + TSG no GPS. Restarted.
|  | 18:00 | [5 empty] | Y |  |  |  | TSG/Matlab 20100906_155040
2010-09-07 | 249 | 6:29 | [9 empty] | ACS frozen, restarted
2010-09-08 | 250 | 8:10 | [9 empty] | ACS frozen, restarted
|  | 15:14 | [9 empty] | ACS frozen, restarted
|  | 16:02 | [9 empty] | ACS frozen, restarted
|  | 18:38 |  |  | on / off | Y |  | Y
|  | 18:42 | OFF | OFF |  |  |  | Y |  |  |  | On Station 67 ctd #9:
|  | 18:50 | [6 empty] | Y |  | Y | New 0.2 um Filter - very dirty
|  | 19:12 | ON | ON | [7 empty] | TSG started
|  | 19:27 | [9 empty] | All OK, some bubbles still in filter. Flow Total/Filtered = 5.1/3.5
2010-09-09 | 251 | 4:06 | [9 empty] | ACS frozen, restarted
|  | 5:55 | [9 empty] | ACS frozen, restarted
|  | 6:36 | [9 empty] | ACS frozen, restarted
|  | 7:52 | [9 empty] | ACS frozen, restarted
|  | 11:27 | [9 empty] | ACS frozen, restarted
|  | 21:35 |  |  |  | Y |  | Y
|  | 23:51 | [9 empty] | FRRF found on "0" reset to "3"
2010-09-10 | 252 | 6:20 | [9 empty] | TSG stopped since 02:00. ACS frozen
|  | 11:20 | [9 empty] | ACS set schedue restarted. UG. Off since 21:35
|  | 18:35 | [9 empty] | TSG lost bluetooth
|  | 18:50 |  |  |  | Y |  | Y
|  | 23:06 | [9 empty] | TSG lost bluetooth + Matlab restarted
2010-09-11 | 253 | 0:15 | [9 empty] | FRRF flow found reversed. 4.6/3.8
|  | 9:50 | OFF | OFF | [7 empty] | Shurflow Pump STOPPED…
|  | 10:24 | [9 empty] | Repaced with spare...Found bypass V2 CLOSED (while in Dry Dock?), & pump overloaded
|  | 10:36 | ON | ON | [7 empty] | All good. FRRF cleaned at intake
|  | 11:11 | [8 empty] | Y | New 0.2 um Filter
2010-09-12 | 254 | 1:17 | [9 empty] | ACS frozen, restarted
|  | 15:56 | [9 empty] | ACS frozen, restarted
|  | 18:29 |  |  |  | Y |  | Y
2010-09-13 | 255 | 7:02 | [9 empty] | ACS frozen, restarted
|  | 7:51 | [9 empty] | ACS frozen, restarted
|  | 19:15 |  |  |  | Y |  | Y
2010-09-14 | 256 | 18:44 |  |  |  | Y |  | Y
2010-09-15 | 257 | 7:50 | [9 empty] | ACS frozen, restarted
|  | 19:00 |  |  |  | Y |  | Y
2010-09-16 | 258 | 5:58 | [9 empty] | ACS frozen, restarted
|  | 8:00 |  |  | on / off | [6 empty] | FRRF bleached 15min soak
|  | 9:00 | [9 empty] | FRRF set to #3
|  | 16:53 | [9 empty] | ACS frozen, restarted
|  | 18:16 |  |  |  | Y |  | Y
2010-09-17 | 259 | 20:50 |  |  |  | Y |  | Y
2010-09-18 | 260 | 7:00 | [9 empty] | ACS frozen, restarted
|  | 15:40 |  |  |  | Y |  | Y
|  | 15:47 | [9 empty] | ACS frozen, restarted
|  | 15:59 | [9 empty] | ACS frozen, restarted
|  | 16:07 | [9 empty] | ACS frozen, restarted
|  | 23:16 | [9 empty] | TSG error - restarted
2010-09-19 | 261 | 1:16 | [9 empty] | TSG date error - restarted
|  | 6:33 | [9 empty] | ACS frozen, restarted
|  | 7:56 | [9 empty] | TSG error - restarted
|  | 7:58 | [9 empty] | ACS frozen, restarted
|  | 8:01 | [9 empty] | ACS losing data like stuttering… cycled power + closed compass
|  | 19:47 |  |  |  | Y |  | Y
2010-09-20 | 262 | 9:26 | [9 empty] | FRRF: in script set transmission = 1; slide = 3
|  | 19:10 |  |  |  | Y |  | Y
2010-09-21 | 263 | 4:43 | [9 empty] | ACS frozen, restarted
|  | 10:51 | [9 empty] | ACS frozen, restarted
|  | 16:05 | [6 empty] | Y |  | Y | New 0.2 um Filter
|  | 16:57 | [9 empty] | ACS frozen, restarted
|  | 17:33 | [9 empty] | ACS frozen, restarted
|  | 20:30 |  |  |  | Y |  | Y
2010-09-22 | 264 | 5:38 | [9 empty] | Power cut to Dry Lab, restarted FRRF + TSG
|  | 13:11 | [9 empty] | ACS frozen, restarted
|  | 14:05 | [9 empty] | ACS frozen, restarted
|  | 17:08 |  |  |  | Y |  | Y
2010-09-23 | 265 | 19:30 | [9 empty] | ACS frozen, restarted
|  | 21:12 |  |  |  | Y |  | Y
2010-09-24 | 266 | 14:00 | [9 empty] | ST HELENA ARRIVAL
|  | 20:59 |  |  |  | Y |  | Y
2010-09-25 | 267 | 8:30 | [9 empty] | Departure from ST HELENA
|  | 18:25 |  |  |  | Y |  | Y |  |  |  | FRR filter left on 0 !
|  | 21:30 | [9 empty] | FRR filter set back on 1 !
2010-09-26 | 268 | 19:58 |  |  |  | Y |  | Y
2010-09-27 | 269 | 3:20 | [9 empty] | ACS frozen, restarted
|  | 18:44 |  |  |  | Y |  | Y
2010-09-28 | 270 | 11:25 | [9 empty] | ACS frozen, restarted
[6 empty] | Y |  | Y
2010-09-29 | 271 | 20:20 |  |  |  | Y |  | Y
2010-09-30 | 272 | 3:01 | [9 empty] | ACS frozen, restarted
|  | 5:00 | [9 empty] | ACS frozen, restarted
|  | 17:00 | [9 empty] | ARRIVAL ASCENSION ISLAND
|  | 17:01 | OFF | OFF | OFF | Y |  | Y | Y | Y | Y | New 0.2 um Filter
2010-10-03 | 275 | 8:00 | [9 empty] | DEPARTURE ASCENSION ISLAND
|  | 8:30 | ON | ON | ON | [6 empty] | All on, all OK
|  | 20:05 |  |  |  | Y |  | Y
2010-10-04 | 276 | 6:58 | [9 empty] | ACS losing records - restarted
|  | 18:48 |  |  |  | Y |  | Y
2010-10-05 | 277 | 0:05 | [9 empty] | FRRF laptop found a virus during SYMANTEC scheduled scan, stopped + rebooted
|  | 23:50 |  |  |  | Y |  | Y
2010-10-06 | 278 | 19:00 |  |  |  | Y |  | Y |  |  |  | Deleted September ACS data from Laptop
|  | 21:18 | [9 empty] | TSG MatLab found OFF
2010-10-07 | 279 | 8:17 | [9 empty] | ACS frozen - restarted
|  | 8:40 | [9 empty] | FLOW LOW: 2.5
|  | 8:50 | INTERMITTANT DATA | [8 empty] | Trying to isolate the suspected blockage… from Temp sensor, OK
|  | 9:30 | " | [8 empty] | FLOW LOW after about 30 minutes
|  | 12:10 | " | [8 empty] | Checking debubbler, ACS, FlowSensor, Filter … FLOW LOW after about 30 minutes
|  | 16:00 | " | [8 empty] | Have bypassed filter + valve, FLOW LOW after about 30 minutes
|  | 17:29 | " |  |  | Y |  | Y |  |  |  | Checked TSG flowthrough, FLOW LOW after about 30 minutes
|  | 18:35 | " | [8 empty] | Back flush with compressed air (tubing only), very dirty water… photos taken
2010-10-08 | 280 | 7:20 | " | [5 empty] | Y |  | Y | New 0.2 um Filter (bypassed). PUMP on again, FLOW LOW after about 30 minutes
|  | 8:00 | " | [8 empty] | Investigating filter in "Sea Chest" box, found clean.
|  | 14:47 | " | [8 empty] | Pump brushes found very worn inside pump. Refurbished and replaced. Pump ON
|  | 15:16 | OFF | OFF | OFF | [6 empty] | PUMP STOPPED - bugger. Juien (engineer) working on a fix.
2010-10-09 | 281 | 21:13 | ON | ON | ON | [6 empty] | After SHORT STN, installed replacement PUMP
|  | 21:42 |  |  |  | Y |  | Y |  |  |  | Total = 5.1, Filtered = 4.4
2010-10-10 | 282 | 7:20 | ON | [8 empty] | ACS set schedule found NOT SET. Total = 4.6, Filtered = 4.1
|  | 20:26 |  |  |  | Y |  | Y
|  | 20:45 | [6 empty] | Y |  |  | ACS cleaned with soapy water, etoh, DI to get rid of any oily residues from pump
2010-10-11 | 283 | 9:10 | OFF | OFF | OFF | [6 empty] | PUMP STOPPED - replacement pump bearings gone
|  | 16:11 |  | ON | ON | [6 empty] | PUMP ON - refurbished with new bearings
|  | 17:01 | ON | [8 empty] | ACS ON - power supply to this restored after work done elsewhere onboard
|  | 20:56 |  |  |  | Y |  | Y
2010-10-12 | 284 | 6:54 | [9 empty] | FRRF found a virus but did not stop
|  | 18:45 | [9 empty] | PUMP OFF for 5 minutes only to re-grease pump
|  | 22:55 |  |  |  | Y |  | Y
2010-10-13 | 285 | 20:03 | OFF | OFF | OFF | [6 empty] | PUMP OFF for maintenance
|  | 21:00 | ON | ON | ON | [6 empty] | PUMP ON
|  | 21:23 |  |  |  | Y |  | Y
2010-10-14 | 286 | 11:15 | OFF | OFF | OFF | [6 empty] | PUMP OFF - on station
|  | 20:58 | ON | ON | ON | [6 empty] | PUMP ON
|  | 21:10 |  |  |  | Y |  | Y
2010-10-15 | 287 | 18:57 | OFF | OFF | OFF | [6 empty] | PUMP OFF - smoking in fact
|  | 22:50 |  |  |  | Y |  | Y
2010-10-16 | 288 | [10 empty] | On station - nothing done
2010-10-17 | 289 | 22:26 |  |  |  | Y |  | Y
2010-10-18 | 290 | 20:50 | [7 empty] | Y
|  | 22:54 |  |  |  | Y |  | Y
2010-10-19 | 291 | 14:06 | [9 empty] | FRRF laptop found a virus during SYMANTEC scheduled scan, stopped + rebooted
|  | 23:01 |  |  |  | Y |  | Y
2010-10-20 | 292 | 12:25 | [9 empty] | FRRF found FAILED, restarted
|  | 20:54 |  |  |  | Y |  | Y
2010-10-21 | 293 | 23:17 |  |  |  | Y |  | Y
2010-10-22 | 294 | 11:01 |  |  |  | Y |  | Y |  |  |  | Cleaned out FRRF manifold
ARRIVAL RIO DE JANEIRO
[11 empty] | Y | New 0.2 um Filter
2010-10-31 | 303 | 15:15 | ON | ON | ON | [6 empty] | All On,All OK
|  | 19:31 | [9 empty] | Changed FRRF settings to position 1
Virus detected on FRRF laptop
2010-11-01 | 304 | 9:00 |  |  |  | Y |  | Y
2010-11-02 | 304 | 10:50 | OFF | OFF | OFF | [6 empty] | All underway system stopped (power failure)
|  | 22:00 | ON | ON | ON | Y |  | Y |  |  |  | All systems ON,
2010-11-03 | 305 | 19:45 |  |  |  | Y
|  | 20:00 | OFF | OFF | OFF | [6 empty] | Power failure
|  | 20:30 | ON | ON | ON
2010-11-04 | 306 |  | ON |  |  | Y |  | Y |  |  |  | ACS frozen
2010-11-05 | 307 |  |  |  |  | Y |  | Y | Y
2010-11-06 | 308 | 10:40 | ON |  |  | Y |  | Y |  |  |  | ACS error "foating point divided by zero"
2010-11-07 | 309 | 3:00 |  |  |  | Y
|  | 21:40 | ON |  |  |  |  | Y |  |  |  | ACS frozen. Discovered that scheduler was not launched _ 5 days of data lost.
2010-11-08 | 310 | 4:00 |  |  |  | Y
|  | 18:00 | [5 empty] | Y
2010-11-09 | 311 | 10:30 | ON |  |  |  |  | Y |  |  |  | ACS frozen
2010-11-10 | 312 |  |  |  |  | y |  | y
|  | 14:55 | ON | [8 empty] | ACS frozen
2010-11-11 | 313 | 3:20 |  |  |  | y |  | y
|  | 11:06 | ON | [8 empty] | ACS frozen
|  | 12:56 | [6 empty] | y
|  | 16:32 | ON | [8 empty] | ACS frozen
2010-11-12 | 314 | 10:29 | ON |  |  | y |  | y |  |  |  | ACS frozen
|  | 13:24 | [9 empty] | Relance Tsg
2010-11-13 |  | 15:17 | OFF | OFF | OFF | [6 empty] | Entrance in Rio de la Plata
ARRIVAL BUENOS AIRES
[11 empty] | Y | New 0.2 um Filter
2010-11-25 | 328 | 9:00 | [9 empty] | LEAVE BUENOS AIRES
2010-11-26 | 329 | 10:55 | ON | ON | [7 empty] | Flow = 5.1 / 4.4. FRRF blockage
|  | 23:16 |  |  | ON | Y | [5 empty] | FRRF bleached. Finally Transmission SET = 2. Block at outflow, for now waste into bucket.
2010-11-27 | 330 | 9:48 | OFF | OFF | OFF | [6 empty] | Found all OFF, no power
|  | 10:09 | ON | ON | ON
|  | 13:45 | [8 empty] | Y | New 0.2 um Filter
|  | 21:35 |  |  |  | Y |  | Y
2010-11-28 | 331 | 14:17 | [6 empty] | Y |  |  | Clean ACS
|  | 23:30 |  |  |  | Y |  | Y
2010-11-29 | 332 | 5:00 | [9 empty] | ACS frozen
|  | 8:44 | [9 empty] | ACS frozen
|  | 12:49 | [9 empty] | ACS frozen
|  | 22:50 |  |  |  | Y |  | Y
2010-11-30 | 333 | 10:41 | OFF | OFF | OFF | [6 empty] | Found POWER OFF, restarted
|  | 10:44 | ON | ON | ON
|  | 22:59 |  |  |  | Y |  | Y |  |  |  | FRRF transmission set =1
2010-12-01 | 334 | 10:12 | [9 empty] | Found POWER off
|  | 13:08 | [9 empty] | Found TSG, FRRF off
|  | 13:54 | [9 empty] | Found TSG, FRRF off
|  | 19:17 | [6 empty] | Y
|  | 23:26 |  |  |  | Y |  | Y
2010-12-02 | 335 | 8:05 | [9 empty] | Found all off, big problems with ship's power all night;
|  | 23:00 | [5 empty] | Y
2010-12-03 | 336 | 18:39 | [9 empty] | Found POWER off
2010-12-04 | 337 | 0:12 |  |  |  | Y |  | Y
|  | 19:03 | OFF | OFF | [5 empty] | Y |  | ALL TURNED OFF 55 - 60 kn !!!! Debubbler waste tubing removed to close deck vent.
|  | 21:10 | [9 empty] | Reset UTC time on all 3 laptops: ACS was + 50sec, TSG + 3mins, FRRF + 1min
|  | 23:28 |  |  |  | Y |  | Y
2010-12-05 | 338 | 18:13 |  |  | OFF | [6 empty] | Found POWER OFF - to FRRF
|  | 18:34 |  |  | ON | [6 empty] | POWER ON
|  | 21:25 | [6 empty] | Y
|  | 22:33 | ON | ON | [7 empty] | Debubbler waste pipe drilled through vent wall….
2010-12-06 | 339 | 23:45 |  |  |  | Y |  | Y
2010-12-07 | 340 | 5:16 | [5 empty] | Y
|  | 19:36 | [9 empty] | Found POWER OFF
|  | 23:29 |  |  |  | Y |  | Y
2010-12-08 | 341 | 13:50 | [9 empty] | Pump off 10 mins - maintenance to Seacock on hull
|  | 15:00 | [9 empty] | Pump off 10 mins - maintenance to Seacock on hull
|  | 19:33 | [6 empty] | Y
|  | 23:52 |  |  |  | Y |  | Y
2010-12-09 | 342 | 1:29 | [9 empty] | ACS frozen
|  | 19:02 | [9 empty] | Found POWER OFF
|  | 19:15 | [8 empty] | Y | New 0.2 um Filter
|  | 22:15 |  |  |  | Y |  | Y |  |  |  | Found debubbler waste not flowing…. Blocked on deck by spare boom
2010-12-10 | 343 | 10:42 | [9 empty] | Found TSG/ACS OFF
2010-12-11 | 344 | 0:25 |  |  |  | Y |  | Y
|  | 10:24 | [9 empty] | Found ACS attenuation NOT GOOD - signal up to 30
|  | 11:00 | [6 empty] | Y |  |  | Cleaned ACS + connector, emailed for further advice…
|  | 18:31 | OFF | OFF |  | Y |  |  |  | Y |  | AT ANCHOR, Magellan Channel
|  | 23:32 |  |  |  | Y
2010-12-12 | 345 | 23:00 |  |  |  | Y |  | Y
2010-12-14 | 347 | 20:00 | [9 empty] | LEAVE ANCHORAGE
|  | 20:10 | ON | ON | [7 empty] | ACS: menu-channels-attenuation-<untick> visible in order to just show absorption data
|  | 22:25 |  |  |  | Y |  | Y
|  | 23:00 | [9 empty] | Found POWER OFF
2010-12-15 | 348 | 3:00 | [9 empty] | Found POWER OFF
|  | 17:26 |  |  |  | Y |  | Y
|  | 19:55 | [9 empty] | Found POWER OFF
|  | 20:20 | [6 empty] | Y |  | Y | New 0.2 um Filter
2010-12-16 | 349 | [10 empty] | Several times found ACS screen frozen…
|  | 22:00 |  |  |  | Y |  | Y
2010-12-17 | 350 | 3:00 | [9 empty] | Power cut
|  | 6:06 | [9 empty] | ACS frozen
|  | 10:44 | [9 empty] | Power CUT OFF
|  | 12:45 | OFF | OFF | [5 empty] | Y |  | AT ANCHOR, Bahia Aguirre
|  | 0:16 |  |  |  | Y |  | Y
2010-12-18 | 351 | 22:00 |  |  |  | Y |  | Y
2010-12-19 | 352 | 9:15 | ON | ON | [7 empty] | LEAVE ANCHORAGE
|  | 21:30 | OFF | OFF | OFF | Y |  | Y | Y | Y |  | ARRIVE USHUAIA. Bleach + FW flush FRRF lines. ACS filter OUT.
2010-12-30 | [10 empty] | Y | new filter on Acs , new GPS on TSG , tests OK
2010-12-31 |  | 23:00 | OFF | ON | ON | [6 empty] | leaving P Williams
2011-01-01 | 1 | 3:10 | ON | ON | ON
|  | 9:20 | ON | ON | ON | [6 empty] | Acs frozen , restarted
|  | 16:35 | ON | ON | ON | [6 empty] | Acs frozen , restarted
2011-01-02 | 2 | 1:00 |  |  |  | Y |  | Y
|  | 6:10 | [9 empty] | TSG restarted
|  | 7:50 | [9 empty] | Acs frozen , restarted
|  | 11:10 | [9 empty] | FRRF restarted
|  | 13:25 | [6 empty] | Y |  |  | Power CUT OFF, stopped Acs TSG for cleaning
|  | 22:00 | [9 empty] | Acs frozen , restarted
2011-01-03 | 3 | 0:00 |  |  |  | Y |  | Y
|  | 9:31 | [9 empty] | TSG restarted (GPS failure)
|  | 10:00 | [9 empty] | FRRF restarted
|  | 19:50 | [9 empty] | Acs frozen , restarted
2011-01-04 | 4 | 2:40 |  |  |  | Y |  | Y
|  | 3:10 |  |  | OFF | [6 empty] | FRRF pump : problems
|  | 12:16 |  |  | ON
|  | 14:55 | [9 empty] | Power CUT OFF
|  | 19:00 | [9 empty] | air vent on FRRF outflow through 3way valve on AC : seems better for pumping
2011-01-05 |  | 1:53 | [9 empty] | sea water flow stopped restarted pump
|  | 2:00 |  |  |  | Y |  | Y
|  | 11:21 | [9 empty] | sea water pump dead since last night, changed for spare, reconditionned the old one ( cleaning brushes)
|  | 12:20 | [9 empty] | Power CUT OFF
|  | 14:00 | [9 empty] | TSG stopped /supply back to 24VDC
|  | 16:32 | [9 empty] | ACS frozen
|  | 17:30 | [9 empty] | Power CUT OFF, ACS+FRRF restarted
2011-01-06 |  | 0:30 |  |  |  | Y |  | Y
2011-01-07 |  | 3:00 |  |  |  | Y |  | Y
2011-01-08 |  | 0:00 |  |  |  | Y |  | Y
|  | 15:00 | [6 empty] | Y |  | Y | ACS filter changed
2011-01-09 |  | 1:00 |  |  |  | Y |  | Y
|  | 9:25 | [9 empty] | Acs frozen
2011-01-10 |  | 0:00 |  |  |  | Y |  | Y
|  | 16:30 | [9 empty] | Acs frozen
|  | 18:27 | [9 empty] | Acs frozen
2011-01-11 |  | 0:00 |  |  |  | Y |  | Y | Y |  |  | found FRRF filter in position "0" since ????
2011-01-12 |  | 1:00 |  |  |  | Y |  | Y
|  | 15:00 | [9 empty] | Power CUT OFF, ACS+FRRF restarted
|  | 22:00 | [9 empty] | Acs frozen
2011-01-13 |  | 3:00 |  |  |  | Y |  | Y
|  | 10:16 | [9 empty] | TSG failure
2011-01-14 |  | 3:00 | OFF |  |  | Y |  | Y |  |  |  | Acs stopped ( anchorage at Petrel cove)
|  | 15:55 |  | OFF | OFF
2011-01-16 |  | 7:30 | ON | ON | ON | Y |  | Y
|  | 18:00 | OFF | OFF | OFF | [6 empty] | GE2 problem , still at anchor in petrel Cove
2011-01-17 |  | 13:00 | ON | ON | ON | [6 empty] | left Petrel Cove , en route to Seymour Isl
|  | 20:15 |  |  |  | Y
2011-01-18 |  | 0:15 | OFF | OFF | OFF | [6 empty] | anchored at Seymour Island
2011-01-19 |  | 20:10 | ON | ON | ON | [6 empty] | left Seymour , back to Antarctic sound/ Brown Bluff
2011-01-20 |  | 13:00 | ON | OFF | OFF | [6 empty] | TSG only to cross Antarctic sound
|  | 20:00 | OFF | OFF | OFF | Y |  | Y |  |  |  | anchored at Petrel cove
2011-01-21 |  | 10:00 | ON | OFF | OFF | [6 empty] | TSG only to cross Antarctic sound
|  | 20:00 | OFF | OFF | OFF |  |  | Y |  |  |  | anchored at Brown Bluff
2011-01-22 |  | 6:16 | ON | ON | ON | [6 empty] | left Brown Bluff, en route to Deception island
|  | 13:00 |  |  |  | Y
2011-01-23 |  | 11:18 | [5 empty] | Y
|  | 13:37 | OFF | OFF | OFF | [6 empty] | arriving Deception Island
2011-01-24 |  | 12:00 | [6 empty] | Y | Y
2011-01-25 |  | 11:40 | ON | ON | ON | [6 empty] | leaving Deception island
2011-01-26 |  | 0:15 |  |  |  | Y |  | Y
2011-01-27 |  | 0:30 |  |  |  | Y |  | Y
2011-01-28 |  | 3:00 |  |  |  | Y |  | Y
|  | 11:30 | OFF | OFF | OFF | Y |  | Y | Y | Y |  | stopped all ,, Acs filter removed, Acs by passed ,all valves closed
arriving P Williams, end of the Antarctic leg
2011-02-03 |  | 6:29 | OFF | ON | 0N | [6 empty] | Departure P Williams, start of the Patagonia VIP leg
|  | 19:56 | OFF | OFF | OFF | Y |  | Y | N | Y
2011-02-04 |  | 10:00 | OFF | ON | 0N
2011-02-05 |  | 0:06 |  |  |  | Y |  | Y | N | Y
|  | 1:07 | [9 empty] | FRR filter found on #3 and set to #1
2011-02-06 |  | 1:20 |  |  |  | Y |  | Y | N | Y
|  | 16:20 | OFF | OFF | OFF | Y |  | Y | N | Y
2011-02-07 |  | 11:00 | OFF | ON | 0N
2011-02-08 |  | 10:08 |  |  |  | Y |  | Y | N | Y
|  | 23:27 | OFF | OFF | OFF | Y |  | Y | N | Y
2011-02-09 |  | 8:00 | OFF | ON | 0N
|  | 17:46 | OFF | OFF | OFF | Y |  | Y | N | Y
2011-02-10 |  | 10:15 | OFF | ON | 0N
|  | 14:50 | OFF | OFF | OFF
|  | 22:30 | OFF | ON | 0N
2011-02-11 |  | 20:18 |  |  |  | Y |  | Y | N | Y |  | Stop FRRF
|  | 21:52 | [9 empty] | Restart FRRF
2011-02-12 |  | 16:20 | OFF | OFF | OFF | Y |  | Y | N | Y
2011-02-13 |  | 7:40 | OFF | ON | 0N
|  | 20:38 | OFF | OFF | OFF | Y |  | Y | N | Y
2011-02-15 |  | 0:30 | OFF | ON | 0N
|  | 13:15 | OFF | OFF | OFF | Y |  | Y | N | Y
2011-02-16 |  | 7:46 | OFF | ON | 0N
|  | 23:10 | OFF | OFF | OFF | Y |  | Y | N | Y |  | Arrival P Montt, end of the Patagonia VIP leg
2011-02-20 |  | 14:30 | OFF | ON | 0N | [6 empty] | Departure P. Montt to Valparaiso
|  | 21:30 | ON | [8 empty] | Start ACS
2011-02-21 |  | 6:30 | [9 empty] | Found TSG stopped since 0:00
|  | 12:08 |  |  |  | Y |  | Y | N | Y
2011-02-22 |  | 2:00 |  |  |  | Y |  | Y | N | Y
|  | 11:05 | [9 empty] | ACS frozen (divide 0)
2011-02-23 |  | 23:30 |  |  |  | Y |  | Y | N | Y
2011-02-24 |  | 20:00 |  |  |  | Y |  | Y | Y | Y
2011-02-25 |  | 15:00 | [9 empty] | PWR out
|  | 16:18 | [9 empty] | Restart FRRF
2011-02-26 |  | 0:40 |  |  |  | Y |  | Y | N | Y
|  | 20:03 | [9 empty] | ACS Frozen
|  | 21:45 | OFF | OFF | OFF | Y |  | Y | Y | Y |  | Stop All, fiter removed, TSG bleached, ACS removed and packed
Arrival Valparaiso
[11 empty] | Y | ACS 091 installed, ACS 007 sent for calibration. Carnet kept here!
2011-03-11 |  | 11:30 | [9 empty] | Depart Valparaiso
|  | 13:30 | ON | 0N | 0N
|  | 22:10 |  |  |  | Y |  | Y |  |  |  | System on, all good. Flow 5.1 / 3.8. FRRF on 3
2011-03-12 |  | 17:45 | [9 empty] | Found TSG failed, restarted
2011-03-13 |  | 1:32 |  |  |  | Y |  | Y
|  | 23:13 |  |  |  | Y |  | Y
2011-03-14 |  | 20:28 | [9 empty] | Flow off for one minute to measure tubing
|  | 21:21 |  |  |  | Y |  | Y |  |  |  | FRRF set = 2
2011-03-15 |  | 14:30 | [9 empty] | Passing ROBINSON CRUSOE ISLAND
|  | 21:55 |  |  |  | Y |  | Y
2011-03-16 |  | 15:30 | [6 empty] | Y |  | Y
|  | 22:50 |  |  |  | Y |  | Y
2011-03-17 |  | 11:41 | [9 empty] | ACS frozen
2011-03-18 |  | 0:02 |  |  |  | Y |  | Y
2011-03-19 |  | 0:56 |  |  |  | Y |  | Y
2011-03-20 |  | 0:02 |  |  |  | Y |  | Y
|  | 12:17 | [9 empty] | adjusted TIME on all laptops, all minor.
|  | 19:12 | [9 empty] | FRRF set=1
|  | 23:56 |  |  |  | Y |  | Y
2011-03-21 |  | 1:14 | [9 empty] | FRRF slide back to 1 after cleaning!
|  | 22:00 | [9 empty] | ThermoSal error in Matlab + failure, just left station 1/2 hour before
|  | 23:56 |  |  |  | Y |  | Y
2011-03-23 |  | 0:41 |  |  |  | Y |  | Y
|  | 12:07 | [9 empty] | ACS found frozen
|  | 23:00 | [6 empty] | Y |  | Y
|  | 23:21 | [9 empty] | All OK!
2011-03-24 |  | 0:35 |  |  |  | Y |  | Y
2011-03-25 |  | 23:28 |  |  |  | Y |  | Y
2011-03-26 |  | 0:35 |  |  |  | Y |  | Y
2011-03-28 |  | 2:04 |  |  |  | Y |  | Y
2011-03-29 |  | 0:53 |  |  |  | Y |  | Y
|  | 11:16 | [9 empty] | Found TSG with stopped with Warning, ACS Floating Point zero error
22:21
2011-03-30 |  | 13:00 | OFF | OFF | OFF |  |  | Y |  | Y |  | Arrival EASTER ISLAND
2011-04-01 |  | 23:00 | [6 empty] | Y |  | Y | Depart EASTER ISLAND
|  | 23:15 | ON | 0N | 0N | [6 empty] | All systems go!
2011-04-02 |  | 14:00 | [9 empty] | ACS freezing
2011-04-02 |  | 15:00 | [9 empty] | Stop start power on ACS to clean connector
2011-04-03 |  | 1:36 |  |  |  | Y |  | Y
2011-04-03 |  | 15:56 | [9 empty] | ACS frozen
2011-04-04 |  | 2:23 |  |  |  | Y |  | Y
2011-04-04 |  | 11:50 | [9 empty] | ACS floating point zero error
2011-04-05 |  | 2:32 |  |  |  | Y |  | Y
2011-04-06 |  | 2:29 |  |  |  | Y |  | Y
2011-04-07 |  | 2:30 |  |  |  | Y |  | Y
|  | 22:08 | OFF | OFF | OFF |  |  |  | Y |  | Y
|  | 22:34 | ON | 0N | 0N
2011-04-08 |  | 3:07 |  |  |  | Y |  | Y |  |  |  | BU on FRRF failed
|  | 14:00 | [9 empty] | FRRF BU
2011-04-09 |  | 3:00 |  |  |  | Y |  | Y
|  | 6:05 | [9 empty] | TSG timed out. Start of period of ship rolling alot…
|  | 13:00 | [9 empty] | TSG timed out, Laptop restarted
2011-04-10 |  | 3:30 |  |  |  | Y |  | Y
|  | 10:30 | [9 empty] | TSG timed out
2011-04-11 |  | 0:36 | [9 empty] | TSG comm port error
|  | 1:32 | [9 empty] | TSG timed out
|  | 2:30 |  |  |  | Y |  | Y |  |  |  | TSG timed out
|  | 3:12 | [9 empty] | TSG timed out
|  | 11:14 | [9 empty] | TSG timed out at 06:07
|  | 12:58 | [9 empty] | TSG timed out
|  | 13:06 | [9 empty] | Power outage FRRF failed
|  | 16:35 | [9 empty] | TSG timed out
2011-04-12 |  | 3:30 |  |  |  | Y |  | Y
|  | 4:31 | [9 empty] | TSG timed out
2011-04-13 |  | 1:40 |  |  |  | Y |  | Y
|  | 11:05 | [9 empty] | TSG timed out
|  | 16:32 | [9 empty] | TSG baud rate errors, restart laptop
2011-04-14 |  | 4:00 |  |  |  | Y |  | Y |  |  |  | Restart FRRF laptop. Took eons, getting old?
|  | 11:25 | [9 empty] | TSG timed out
2011-04-15 |  | 2:46 |  |  |  | Y |  | Y
2011-04-16 |  | 4:45 |  |  |  | Y |  | Y
|  | 6:33 | [9 empty] | FRRF slide found on 0, reset to 1
|  | 10:00 | [9 empty] | TSG timed out
|  | 10:49 | [9 empty] | TSG timed out
2011-04-17 |  | 11:41 | [5 empty] | Y |  |  |  | TSG timed out at 06:35
|  | 22:36 |  |  |  | Y |  | Y
2011-04-18 |  | 5:04 | [9 empty] | TSG timed out
|  | 10:39 | [9 empty] | TSG timed out
|  | 11:58 | [9 empty] | Pwr out to DryLab, FRRF faied, TSG and ACS OK
|  | 15:00 | OFF | OFF | OFF |  |  |  | Y |  | Y
|  | 15:36 | ON | 0N | 0N
2011-04-19 |  | 0:57 |  |  |  | Y |  | Y
|  | 9:07 | [9 empty] | TSG timed out
2011-04-20 |  | 5:16 | [9 empty] | TSG timed out at 02:00
|  | 10:51 | [9 empty] | TSG timed out at 07:05
|  | 12:05 |  |  |  | Y |  | Y
|  | 19:42 | [9 empty] | TSG timed out
2011-04-21 |  | 1:34 |  |  |  | Y |  | Y
|  | 12:12 | [9 empty] | TSG timed out
2011-04-22 |  | 4:00 |  |  |  | Y |  | Y
|  | 5:00 | [9 empty] | Found flow reversed, righted
2011-04-23 |  | 2:18 |  |  |  | Y |  | Y
|  | 6:02 | [9 empty] | Found air leak in FRRF tubing. OK now
|  | 14:20 | [9 empty] | TSG timed out at 09:05
2011-04-24 |  | 7:07 | [9 empty] | TSG timed out
|  | 9:05 | [9 empty] | TSG timed out
|  | 9:26 |  |  |  | Y |  | Y
|  | 13:00 | [9 empty] | TSG timed out
|  | 20:08 | [9 empty] | Found ACS attenuation c-side FAILED (no lights)
2011-04-25 |  | 0:54 | OFF | OFF | OFF | Y |  | Y | Y | Y
Arrival GUAYAQUIL
2011-04-29 |  | 13:00 | [8 empty] | Y | Depart GUAYAQUIL
2011-05-01 |  | 11:40 | ON | 0N | 0N | [6 empty] | FRRF on "2"
|  | 14:52 | [9 empty] | TIME adjusted on Laptops to UTC. <1 minute out.
|  | 23:32 | OFF | OFF | OFF | [6 empty] | AC problem in DryLab, pump turned OFF
|  | 23:49 |  |  |  | Y |  | Y |  |  |  | Changed TSG program : obj1.timeout = 100 (from 50)
2011-05-02 |  | 0:33 | ON | 0N | 0N
|  | 8:37 | [9 empty] | Found TSG failed: i/p buffer full before terminator reached. Index longg(1) out of bounds
2011-05-03 |  | 0:28 |  |  |  | Y |  | Y
2011-05-04 |  | 3:12 |  |  |  | Y |  | Y
|  | 22:49 | [9 empty] | Found TSG failed: ERROR using vert cat. Out of memory.
2011-05-05 |  | 6:02 |  |  |  | Y |  | Y
|  | 14:13 | [9 empty] | Found FRRF FAILED. Restarted
|  | 16:26 | [9 empty] | TSG timed out
|  | 22:54 |  |  |  | Y |  | Y
2011-05-06 |  | 12:31 | [9 empty] | Found FRRF failed, power outage?
2011-05-07 |  | 0:41 |  |  |  | Y |  | Y
|  | 12:35 | OFF | OFF | OFF |  |  |  |  | Y |  | Arrival GALAPAGOS
2011-05-09 |  | 13:50 | ON | 0N | 0N | [5 empty] | Y | Depart GALAPAGOS
2011-05-10 |  | 1:04 |  |  |  | Y |  | Y
|  | 14:07 | [9 empty] | FRRF failed, power outage
2011-05-11 |  | 0:13 |  |  |  | Y |  | Y
2011-05-12
2011-05-13 |  | 5:11 |  |  |  | Y |  | Y
|  | 12:23 | [9 empty] | FRRF water blockage in tubing, backflushed + OK
2011-05-14 |  | 0:34 |  |  |  | Y |  | Y
|  | 23:50 | [9 empty] | Reset TIME on all to GPS UTC
2011-05-15 |  | 0:09 | [9 empty] | TSG error: vertcat CAT argument not consistent. Restarted
2011-05-16 |  | 0:15 |  |  |  | Y |  | Y
|  | 22:57 |  |  |  | Y |  | Y
|  | 11:04 | OFF | OFF | OFF | Y |  | Y | Y | Y
|  | 21:30 | [9 empty] | Arrival GUAYAQUIL
2011-05-19 |  | 23:20 | ON | 0N | 0N | [5 empty] | Y | Depart GUAYAQUIL ,Fluorometer instaled
2011-05-20 |  | 1:20 | [9 empty] | Error TSG
|  | 11:30 | [9 empty] | ACS frozen
|  | 12:48 | [9 empty] | ACS frozen
2011-05-21 |  | 2:00 |  |  |  | Y |  | Y
2011-05-22 |  | 2:24 |  |  |  | Y |  | Y
|  | 11:15 | [9 empty] | TSG cleaned
2011-05-23 |  | 1:37 |  |  |  | Y |  | Y
|  | 8:35 | [9 empty] | Error TSG
2011-05-24 |  | 2:35 |  |  |  | Y |  | Y
|  | 5:10 | [9 empty] | ACS frozen
2011-05-25 |  | 2:16 |  |  |  | Y |  | Y
|  | 23:25 | [6 empty] | Y |  | Y | Stop for cleaning / Back up all
2011-05-26 |  | 2:40 | [9 empty] | Pump changed / TSG data deleted
2011-05-27 |  | 1:05 |  |  |  | Y |  | Y
|  | 2:05 | [9 empty] | FRR on #3
|  | 12:50 | [9 empty] | ACS frozen
2011-05-28 |  | 1:00 |  |  |  | Y |  | Y
|  | 1:30 | [9 empty] | Valve control: was it in filtered since we changed filter?
|  | 12:42 | [9 empty] | ACS frozen
|  | 14:00 | [9 empty] | TSG time out, restart
2011-05-29 |  | 3:35 |  |  |  | Y |  | Y
2011-05-30 |  | 2:57 | [9 empty] | ACS frozen
|  | 3:05 |  |  |  | Y |  | Y
|  | 12:37 | [9 empty] | Error message: FRR's data not saved
2011-05-31 |  | 12:20 |  |  |  | Y |  | Y
2011-06-01 |  | 2:55 |  |  |  | Y |  | Y
|  | 13:08 | [9 empty] | ACS frozen
|  | 15:20 |  |  |  |  | Y |  | Y |  | Y | Fluorometer cleaned, restart all
|  | 17:40 | [9 empty] | Problem TSG matlab
2011-06-02 |  | 0:57 | [9 empty] | ACS frozen
|  | 2:30 |  |  |  | Y |  | Y
|  | 4:32 | [9 empty] | ACS frozen
2011-06-03 |  | 1:20 |  |  |  | Y |  | Y
|  | 16:38 | [9 empty] | FRR: floating point overflow, restart
|  | 16:55 | [9 empty] | FRR: floating point division by zero, stop
|  | 18:44 | [9 empty] | FRR restarted
|  | 23:20 |  |  |  |  | Y |  |  |  |  | Fluorometer cleaned, restart ACS and TSG
2011-06-04 |  | 2:15 |  |  |  | Y |  | Y
|  | 16:35 | [9 empty] | Fuorometer data now OK (Temperature was acquired before!)
|  | 19:30 | [9 empty] | FRR floating point dividion by zero, restart
|  | 20:25 | [9 empty] | FRR: Set filter to position 1
2011-06-05 |  | 0:10 |  |  |  | Y |  | Y
2011-06-06 |  | 0:40 |  |  |  | Y | Y | Y
|  | 13:20 | [9 empty] | ACS frozen
2011-06-07 |  | 0:10 |  |  |  | Y | Y | Y
|  | 0:36 | [9 empty] | Problem after stop, restart again
2011-06-08 |  | 1:00 |  |  |  | Y | Y | Y
2011-06-09 |  | 1:10 |  |  |  | Y | Y | Y
|  | 4:48 | [9 empty] | TSG stopped time out
2011-06-10 |  | 1:00 |  |  |  | Y | Y | Y
|  | 10:00 | [9 empty] | ACS frozen
|  | 11:56 | [9 empty] | TSG problem for 50'
2011-06-11 |  | 1:00 |  |  |  | Y | Y | Y
|  | 3:30 | [9 empty] | ACS frozen, floating point division by zero
|  | 17:10 | [9 empty] | ACS frozen
|  | 21:25 | [9 empty] | AC off since ? Restart FRR
2011-06-12 |  | 2:15 |  |  |  | Y | Y | Y
2011-06-13 |  | 2:00 |  |  |  | Y | Y | Y
|  | 16:11 | [9 empty] | ACS frozen
|  | 18:45 | [9 empty] | Stop all since FW
2011-06-14 |  | 2:00 |  |  |  | Y | Y | Y | Y |  | Y
|  | 5:47 | [9 empty] | ACS frozen
|  | 22:37 | [9 empty] | ACS frozen
2011-06-15 |  | 7:15 |  |  |  | Y | Y | Y
|  | 16:00 | [9 empty] | ACS frozen
2011-06-16 |  | 4:45 |  |  |  | Y | Y | Y
2011-06-17 |  | 2:00 | [9 empty] | Filter found pushed to intermediate position
|  | 2:10 |  |  |  | Y | Y | Y
2011-06-18 |  | 2:00
|  | 16:45 | OFF | OFF | OFF | Y | Y | Y | Y | Y |  | Computer time set, clean all, bleach FRR tubing and intake
Arrival GAMBIER
[7 empty] | Y |  |  |  | Y | START CORAL LEG in Gambiers Lagoon
2011-06-26 |  | 16:00 | OFF | OFF | 0N | [6 empty] | FFRF ONLY ON - pump failed replaced and requested replacement spare
2011-06-28 |  | 3:50 |  |  |  | Y |  | Y |  |  |  | Cuvette full of material
2011-06-29 |  | 3:10 |  |  |  | Y |  | Y
|  | 17:39 | OFF | 0N | 0N | [6 empty] | Bypassing ACS + Fluorometer. Flow = 4.0 - 4.5 l/min
2011-06-30 |  | 3:38 | [9 empty] | TSG timeout error
|  | 4:00 |  |  |  | Y |  | Y
|  | 4:05 | [9 empty] | TSG back ON
|  | 16:44 | [9 empty] | TSG timeout error
2011-07-01 |  | 0:37 | OFF | OFF | OFF | [6 empty] | Found FLOW =0
|  | 2:30 |  | 0N | 0N | [6 empty] | Found Medusa stuck in intake at hull… OK now
|  | 3:38 |  |  |  | Y |  | Y
2011-07-02 |  | 21:10 | [9 empty] | TSG timeout error
|  | 2:45 |  |  |  | Y |  | Y
|  | 16:44 | OFF | OFF | OFF | Y |  | Y |  | Y |  | Alongside at RIKITEA
2011-07-05 |  | 15:59 |  | 0N | 0N | [6 empty] | Depart RIKITEA
2011-07-06 |  | 5:44 |  |  |  | Y |  | Y
|  | 22:58 | [9 empty] | TSG timeout error
2011-07-07 |  | 2:58 |  |  |  | Y |  | Y
|  | 16:20 | [9 empty] | TSG timeout error
2011-07-08 |  | 3:25 |  |  |  | Y |  | Y
2011-07-09 |  | 2:16 |  |  |  | Y |  | Y
2011-07-10 |  | 2:21 |  |  |  | Y |  | Y
|  | 16:00 | [9 empty] | Found TSG laptop asleep!! Having stopped at 05:02
|  | 18:04 | [9 empty] | TSG finally working now
2011-07-11 |  | 3:19 |  |  |  | Y |  | Y
|  | 17:57 | OFF | OFF | OFF | Y | Y | Y |  | Y |  | END OF CORAL LEG - alongside at RIKITEA
Replaced FRRF tubing
2011-07-16 |  | 14:30 | [8 empty] | Y | Depart RIKITEA
|  | 15:56 | 0N | 0N | 0N
2011-07-17 |  | 9:30 |  |  |  | Y |  | Y
|  | 16:09 | [9 empty] | Found FRRF failed
2011-07-18 |  | 4:00 |  |  |  | Y |  | Y
2011-07-19 |  | 3:09 | [9 empty] | ACS very FROZEN
|  | 5:20 |  |  |  | Y |  | Y
|  | 17:39 | [9 empty] | Found FRRF failed
2011-07-20 | [11 empty] | FOUND VALVE ON WRONG SETTING: 30mins filtered 5mins Total
|  | 6:00 |  |  |  | Y |  | Y
|  | 15:05 | [9 empty] | TSG error, had stopped at 14:33, restarted WITHOUT Flour
2011-07-21 |  | 3:30 | OFF | OFF | OFF | Y |  | Y |  | Y |  | Arrive HIVA OA
2011-07-22 |  | 6:10 | 0N | 0N | 0N | [6 empty] | Depart HIVA OA, Fluor back ON, OK
|  | 20:30 | OFF | OFF | OFF | Y | Y | Y | Y | Y |  | Arrive NUKU HIVA
2011-07-26 |  | 2:45 | 0N | 0N | 0N | [6 empty] | Depart NUKU HIVA (anchorage in north)
|  | 2:46 | [9 empty] | Fluorometer Canister LEAKING
|  | 3:03 | [9 empty] | TSG online minus Fluor
2011-07-27 |  | 4:00 |  |  |  | Y |  | Y |  |  |  | Fluor back inline and NOT LEAKING (sealed with silicone)
|  | 15:52 | [9 empty] | TSG error, restarted
2011-07-28 |  | 4:20 |  |  |  | Y |  | Y
2011-07-29 |  | 9:50 |  |  |  | Y |  | Y
|  | 12:00 | [9 empty] | Arrive anchorage NUKU HIVA
|  | 17:50 | OFF | OFF | OFF |  |  | Y
[6 empty] | Y |  | Y | Y | Y | Y
2011-07-31 |  | 16:40 | 0N | 0N | 0N | [6 empty] | Depart anchorage NUKU HIVA
2011-08-01 |  | 5:15 |  |  |  | Y |  | Y
|  | 17:20 | [9 empty] | FRRF failed - restarted
2011-08-02 |  | 2:30 | OFF | OFF | OFF | Y | Y | Y | Y | Y
Arrive anchorage NUKU HIVA (west)
2011-08-04 |  | 2:15 | 0N | 0N | 0N | [6 empty] | Depart anchorage NUKU HIVA (west)
|  | 16:43 | OFF | [8 empty] | Stopped ACS to read Fluor for station position, ie no filtered interruptions
|  | 20:00 | 0N
2011-08-05 |  | 1:40 |  |  |  | Y |  | Y
2011-08-06 |  | 8:15 | OFF | OFF | OFF | Y |  | Y |  |  |  | Arrive anchorage UA POA
|  | 17:00 |  |  |  | Y |  | Y
|  | 18:00 | [6 empty] | Y | Y | Y
2011-08-07 |  | 9:10 | 0N | 0N | 0N | [6 empty] | Depart anchorage UA POA
2011-08-08 |  | 1:00 |  |  |  | Y |  | Y
|  | 3:00 | [9 empty] | TSG timeout error - restarted
|  | 15:19 | [9 empty] | TSG timeout error - restarted
|  | 16:00 | [9 empty] | TSG timeout error - restarted
|  | 16:27 | [9 empty] | TSG timeout error - restarted
|  | 16:36 | [9 empty] | … working OK without Fluor (in MatLab script)
2011-08-09 |  | 18:50 |  |  |  | Y |  | Y
2011-08-10 |  | 2:46 | [9 empty] | …working OK again WITH Fluor (in MatLab script), after checking all connectors
2011-08-11 |  | 9:40 | [9 empty] | TSG timeout error - restarted
|  | 18:50 |  |  |  | Y |  | Y
|  | 16:15 | OFF | OFF | OFF | [6 empty] | Arrive anchorage FAKAREVA
2011-08-15 |  | 1:00 | 0N | 0N | 0N | Y | [5 empty] | Depart anchorage FAKAREVA
|  | 19:50 | [9 empty] | ACS frozen - screen showed no data
|  | 23:33 | [9 empty] | ACS stopped
2011-08-16 |  | 6:05 | [9 empty] | ACS stopped
|  | 16:10 | OFF | OFF | OFF |  |  | Y |  | Y |  | Arrive PAPEETE
[6 empty] | Y | Y |  | Y |  | Y
2011-08-26 |  | 20:45 | 0N | 0N | 0N | [6 empty] | Depart MOOREA (TAHITI)
2011-08-27 |  | 1:10 | [9 empty] | TSG error - changed USB hub and restarted 01:40
|  | 5:09 |  |  |  | Y |  | Y
|  | 6:57 | [9 empty] | TSG error
|  | 15:45 | [9 empty] | TSG restarted
|  | 18:30 | [9 empty] | Updated all dry lab laptops with correct UTC time
2011-08-28 |  | 2:31 | [9 empty] | TSG timeout error - restarted
|  | 3:00 |  |  |  | Y |  | Y |  |  |  | (FRRF cleaning interrupted by muster alarm)
|  | 15:48 | [9 empty] | TSG stopped, restarted
2011-08-29 |  | 6:53 |  |  |  | Y |  | Y |  |  |  | TSG hub issues - reinstalled driver software, reboot at 07:18
|  | 11:01 | [9 empty] | TSG error, ACS frozen
|  | 16:15 | [9 empty] | TSG timeout ~ 11:37!
2011-08-30 |  | 5:03 |  |  |  | Y |  | Y |  |  |  | Note - FRRF found in state "0" instead of "1" - potentially in that state since 06:53 previous day?
|  | 12:24 | [9 empty] | TSG timeout ~ 0805
2011-08-31 |  | 2:40 |  |  |  | Y |  | Y
|  | 11:43 | [9 empty] | TSG timeout @ 0701
|  | 12:30 | [9 empty] | TSG timeout @ 1212
|  | 14:11 | [9 empty] | TSG baud rate error @1326
2011-09-01 |  | 2:43 |  |  |  | Y |  | Y
|  | 4:30 | [9 empty] | Found FRRF failed
|  | 8:02 | [9 empty] | TSG timeout @ 0518
2011-09-02 |  | 3:07 |  |  |  | Y |  | Y
|  | 16:47 | [9 empty] | TSG timeout
2011-09-03 |  | 3:13 |  |  |  | Y |  | Y
|  | 4:33 | [9 empty] | Found FRRF failed
|  | 8:13 | [9 empty] | TSG timeout
|  | 17:06 | [9 empty] | Comm port error TSG
|  | 18:17 | [9 empty] | TSG timeout error 1725. Found FRRF pump switched off and disconnected after Accuri sample
2011-09-04 |  | 0:40 | [9 empty] | TSG timeout
|  | 2:16 | [9 empty] | TSG timeout
|  | 4:05 | OFF | OFF | OFF |  |  |  | Y | Y | Y | STOP SYSTEM, filter clean ACS/FLOUR. RESTART SYSTEM 0428
|  | 4:35 | 0N | 0N | 0N | Y |  | Y
|  | 9:27 | [9 empty] | TSG comm error
|  | 11:40 | [9 empty] | Found ACS frozen
|  | 15:55 | [9 empty] | TSG timeout @ 1237
|  | 19:52 | [9 empty] | TSG timeout @ 1842
|  | 21:57 | [9 empty] | TSG timeout @ 2013
2011-09-05 |  | 4:45 |  |  |  | Y |  | Y |  |  |  | ACS laptop crashed - reboot @ 0458
|  | 5:05 | [9 empty] | TSG timeout
|  | 12:43 | [9 empty] | TSG timeout @ 1230
|  | 15:38 | [9 empty] | TSG timeout @ 1518
|  | 19:25 | [9 empty] | TSG timeout @ 1644
|  | 22:30 | [9 empty] | TSG timeout @ 1936
2011-09-06 |  | 4:15 | [9 empty] | TSG timeout @ 0400
|  | 5:09 |  |  |  | Y |  | Y
|  | 16:19 | [9 empty] | FRRF failed
2011-09-07 |  | 2:40 | [9 empty] | FRRF failed
|  | 4:07 | [9 empty] | TSG timeout @ 0338
|  | 4:46 | OFF | OFF | OFF |  |  |  | Y | Y |  | Stopping underway, freshwater clean. KIRIBATI
2011-09-08 |  | 5:30 | 0N | 0N | 0N | [6 empty] | Restart systems after stopover in Kiribati
2011-09-09 |  | 5:00 |  |  |  | Y |  | Y
|  | 21:34 | [9 empty] | TSG timeout aspirator(?) 21:22
2011-09-10 |  | 3:30 |  |  |  | Y |  | Y
|  | 15:48 | [9 empty] | ACS frozen/FRRF failed
2011-09-11 |  | 3:42 |  |  |  | Y |  | Y
|  | 10:40 | [9 empty] | TSG timeout 0401
|  | 12:10 | [9 empty] | "Warning! Input buffer filled before terminator reached" - issue revolves around allocation of serial port in matlab. Need to alter in device manager, then RESTART THE COMPUTER, then run script in matlab as per usual.
|  | 19:00 | [9 empty] | Restart TSG
2011-09-12 |  | 0:04 |  |  |  | Y |  | Y
|  | 16:09 | [9 empty] | TSG timeout @ 0812
|  | 16:51 | [9 empty] | FRRF slide 1/2 in
TSG timeout @ 1636
2011-09-13 |  | 3:35 |  |  |  | Y |  | Y
|  | 12:33 | [9 empty] | TSG timeout @ 1123
2011-09-14 |  | 0:57 |  |  |  | Y |  | Y
|  | 17:30 | [9 empty] | TSG timeout
|  | 22:30 | OFF | OFF | OFF | [5 empty] | Y | Stopped TSG and ACS for filter cleaning
|  | 23:09 | 0N | 0N | 0N | [6 empty] | Restart system
2011-09-15 |  | 3:24 |  |  |  | Y |  | Y
|  | 12:14 | [9 empty] | ACS crashed, FRRF crashed with bad date stamp
2011-09-16 |  | 4:43 |  |  |  | Y |  | Y
|  | 12:11 | [9 empty] | TSG timeout @ 07:04
|  | 13:03 | [9 empty] | TSG timeout @ 12:28
2011-09-17 |  | 2:38 |  |  |  | Y |  | Y
|  | 19:48 | [9 empty] | TSG timeout @ 17:20, FRRF crashed, ACS crashed….. Power issue?
2011-09-18 |  | 0:39 |  |  |  | Y |  | Y
|  | 20:43 | OFF | OFF | OFF | Y | Y | Y | Y | Y |  | Arrive HONOLULU
2011-09-28 |  | 21:09 | 0N | 0N | OFF | [5 empty] | Y | Depart HONOLULU 'HAWAI) pb with new computer for FRR- password missing
2011-09-29 |  | 1:14 |  |  | 0N | [6 empty] | start FRR with old computer
|  | 5:00 |  |  |  | Y |  | Y
|  | 20:12 | [9 empty] | TSG Timeout
2011-09-30 |  | 4:25 |  |  |  | Y |  | Y
|  | 6:28 | [9 empty] | TSG Timeout
|  | 16:25 | [9 empty] | TSG Timeout
|  | 19:49 | [9 empty] | FRR failed
2011-10-01 |  | 4:25 |  |  |  | Y |  | Y |  |  |  | forget to start FRR (claudie)
|  | 6:14 | [9 empty] | Start FRR
|  | 8:02 | [9 empty] | TSG Timeout
2011-10-02 |  | 3:40 | [9 empty] | ACS warning (error floating point…)
|  | 3:56 |  |  |  | Y |  | Y
|  | 8:53 |  |  | OFF | [6 empty] | FRR xomputer stopped, power box out, new power suply, ok
0N
2011-10-03 |  | 1:02 | [9 empty] | TSG Timeout
|  | 4:36 |  |  |  | Y |  | Y |  |  |  | change FRR computer (new one)
|  | 9:25 | [9 empty] | ACS warning (error floating point…)
2011-10-04 |  | 4h28 |  |  |  | Y |  | Y |  |  |  | FRR computer was not set at right time 09h25pm on 3rd october in spite of 4h29 on 4th october
|  | 19h58 | [9 empty] | files acquired with a bad date are stored in a new folder ünderway/wrong date (files 275 and 276)
2011-10-05 | [11 empty] | ACS frozen
|  | 04h12 |  |  |  | Y |  | Y |  |  |  | TSG timeout @ 01:53
|  | 16h50 | [9 empty] | FRR failed to attempt to open driver in FIFO command pipe 2
|  | 22h55 | [9 empty] | 2 errors:floating point division by 0
2011-10-06 |  | 03h45 |  |  |  | Y |  | Y | Y |  | Y
|  | 05h15 | OFF | OFF | OFF | [6 empty] | POWER outage supply 220v & 24Vdc all stopped
|  | 05h20 | 0N | 0N | 0N | [6 empty] | all re-started
|  | 6:30 | [9 empty] | POWER outage- restart ACS- FRR
|  | 6:32 | OFF | OFF | OFF | [6 empty] | POWER outage ACS-TSG-FRR
|  | 7:17 | 0N | 0N | 0N | [6 empty] | restart ACS & FRR
|  | 19:51 | OFF | OFF | OFF | [6 empty] | stop all- changeACS filter
|  | 21h05 | 0N | 0N | 0N | [6 empty] | restart all
2011-10-07 |  | 4:33 |  |  |  | Y |  | Y |  |  |  | during FRR clean up I forgot (Claudie)to stop its acquisition but filter well positionned
2011-10-08 |  | 4h05 |  |  |  | Y |  | Y
2011-10-09 |  | 4:08 |  |  |  | Y |  | Y
2011-10-10 |  | 4:00 |  |  |  | Y |  | Y |  |  |  | TSG timeout @ 17:40 for 1h30
2011-10-11 |  | 4:16 |  |  |  | Y |  | Y
2011-10-12 |  | 04;35 |  |  |  | Y |  | Y |  |  |  | change FRR subing of pump
2011-10-13 |  | 4:17 |  |  |  | Y |  | Y
2011-10-14 |  | 3:30 |  |  |  | Y | Y | Y | Y |  | Y | set FRR computer time
2011-10-15 |  | 3:20 |  |  |  | Y |  | Y
2011-10-16 |  | 2:27 |  |  |  | Y |  | Y
2011-10-17 |  | 2:15 |  |  |  | Y |  | Y
2011-10-18 |  | 2:42 |  |  |  | Y |  | Y |  |  |  | FRR pump found out of order (since when?) we change pump
2011-10-19 |  | 1:50 |  |  |  | Y |  | Y
|  | 15:19 | [9 empty] | TSG Matlab out of memory
2011-10-20 |  | 1:35 |  |  |  | Y |  | Y
2011-10-21 |  | 3:29 |  |  |  | Y |  | Y
|  | 15:05 | [9 empty] | ACS frozen
|  | 20:10 | [9 empty] | ACS frozen
2011-10-22 |  | 0:25 | [9 empty] | ACS frozen
|  | 3:57 |  |  |  | Y |  | Y
2011-10-23 |  | 1:35 |  |  |  | Y |  | Y
2011-10-24 |  | 1:50 |  |  |  | Y |  | Y
|  | 13:11 | [9 empty] | TSG time out for '4 hours
|  | 17:15 |  |  |  | Y |  | Y | Y |  | Y | stop TSG and ACS before filter change
2011-10-25 |  | 1:02 |  |  |  | Y |  | Y
2011-10-26 |  | 1:32 |  |  |  | Y |  | Y
|  | 5:00 | [9 empty] | Drive full on ACS laptop
|  | 20:43 | OFF | OFF | OFF | Y | Y | Y | Y | Y |  | Arrive SAN DIEGO
[9 empty] | Y | Y | Y | Depart SAN DIEGO. ACS 091
2011-11-24 |  | 16:30 | 0N | 0N | 0N | [6 empty] | Full of bubbles. Flow 4.1 & 3.4. ACS 560 wavelngth "zero"
|  | 20:35 | [9 empty] | Free of bubbles now…
2011-11-25 |  | 2:40 |  |  |  | Y |  | Y |  |  |  | FRRF Laptop time 1 hour behind UTC, + 1 day ahead. Changed here
2011-11-26 |  | 2:12 |  |  |  | Y |  | Y
|  | 10:02 | [9 empty] | TSG timeout error
2011-11-27 |  | 1:30 | [9 empty] | ACS "floating point zero" error
|  | 2:10 |  |  |  | Y |  | Y
2011-11-28 |  | 3:12 |  |  |  | Y |  | Y
2011-11-29 |  | 5:22 |  |  |  | Y |  | Y
2011-11-30 |  | 3:08 |  |  |  | Y |  | Y
2011-12-01 |  | 1:31 | [9 empty] | TSG timeout error, stopped at 00:01:22, restarted here
|  | 12:11 | OFF | OFF | [7 empty] | Found Flow = 0, it had stopped at 11:51
|  | 12:48 | [9 empty] | New Pump installed, not sounding great, so install refurbed one
|  | 13:28 | 0N | 0N | [7 empty] | ALL GOOD and going again
|  | 16:05 | [9 empty] | ACS "floating point zero" errors
2011-12-03 |  | 3:57 |  |  |  | Y |  | Y
2011-12-04 |  | 6:53 |  |  |  | Y |  | Y
|  | 9:45 | [9 empty] | TSG vertcat error
|  | 17:07 | [9 empty] | TSG timeout
2011-12-05 |  | 0:50 | OFF | OFF |  |  | Y |  | Y |  | Y
|  | 1:30 | 0N | 0N | [7 empty] | Flow 5.6/5.1
|  | 1:40 |  |  |  | Y |  | Y
2011-12-06 |  | 3:24 |  |  |  | Y |  | Y
|  | 7:03 | [9 empty] | ACS "Floating Point" errors
|  | 18:28 | [9 empty] | restarted TSG, seemed very slow
2011-12-07 |  | 3:46 |  |  |  | Y |  | Y
|  | 7:07 | [9 empty] | TSG vertcat error
|  | 13:17 | [9 empty] | ACS frozen
|  | 16:07 | OFF | OFF | OFF |  |  |  |  | Y |  | At Anchor CLIPPERTON
|  | 22:00 | 0N | 0N | 0N | [6 empty] | Depart CLIPPERTON
2011-12-08 |  | 1:26 |  |  |  | Y |  | Y
2011-12-09 |  | 3:48 |  |  |  | Y |  | Y
2011-12-10 |  | 3:15 |  |  |  | Y |  | Y
|  | 8:26 | [9 empty] | FRRF FAILED: restarted
2011-12-12 |  | 2:59 |  |  |  | Y |  | Y
|  | 21:20 | OFF | OFF |  |  | Y |  | Y |  | Y
|  | 21:54 | 0N | 0N
|  | 23:30 | [9 empty] | Bubbles now out?
2011-12-13 |  | 2:20 |  |  |  | Y |  | Y
2011-12-14 |  | 3:59 |  |  |  | Y |  | Y
2011-12-15 |  | 3:03 |  |  |  | Y |  | Y
2011-12-16 |  | 8:47 | [9 empty] | TSG failed- Restarted
|  | 17:03 | OFF | OFF |  |  | Y |  | Y |  | Y
|  | 17:36 | 0N | 0N
2011-12-17 |  | 2:34 |  |  |  | Y |  |  | Y
|  | 18:29 | [9 empty] | ACS absorption side LIGHT FAILURE
2011-12-18 |  | 1:27 |  |  |  | Y |  |  | Y
|  | 13:15 | [9 empty] | At Anchor COCOS ISLANDS
|  | 18:48 | [9 empty] | Depart COCOS ISLANDS
2011-12-19 |  | 3:50 |  |  |  | Y |  |  | Y
2011-12-20 |  | 1:09 | [9 empty] | TSG vertcat error - restarted
|  | 2:26 |  |  |  | Y |  |  | Y
|  | 15:58 | [9 empty] | TSG timeout
2011-12-21 |  | 3:59 |  |  |  | Y |  |  | Y
|  | 6:14 | [9 empty] | TSG vertcat error - restarted
2011-12-22 |  | 11:34 | [9 empty] | TSG vertcat error - restarted
|  | 13:33 | OFF | OFF | OFF | Y | Y | Y | Y | Y |  | At Anchor PANAMA
Y

## Jday
Date |  | jday
2010-01-01 | 40179 | 0
2010-01-02 | 40180 | 1
2010-01-03 | 40181 | 2
2010-01-04 | 40182 | 3
2010-01-05 | 40183 | 4
2010-01-06 | 40184 | 5
2010-01-07 | 40185 | 6
2010-01-08 | 40186 | 7
2010-01-09 | 40187 | 8
2010-01-10 | 40188 | 9
2010-01-11 | 40189 | 10
2010-01-12 | 40190 | 11
2010-01-13 | 40191 | 12
2010-01-14 | 40192 | 13
2010-01-15 | 40193 | 14
2010-01-16 | 40194 | 15
2010-01-17 | 40195 | 16
2010-01-18 | 40196 | 17
2010-01-19 | 40197 | 18
2010-01-20 | 40198 | 19
2010-01-21 | 40199 | 20
2010-01-22 | 40200 | 21
2010-01-23 | 40201 | 22
2010-01-24 | 40202 | 23
2010-01-25 | 40203 | 24
2010-01-26 | 40204 | 25
2010-01-27 | 40205 | 26
2010-01-28 | 40206 | 27
2010-01-29 | 40207 | 28
2010-01-30 | 40208 | 29
2010-01-31 | 40209 | 30
2010-02-01 | 40210 | 31
2010-02-02 | 40211 | 32
2010-02-03 | 40212 | 33
2010-02-04 | 40213 | 34
2010-02-05 | 40214 | 35
2010-02-06 | 40215 | 36
2010-02-07 | 40216 | 37
2010-02-08 | 40217 | 38
2010-02-09 | 40218 | 39
2010-02-10 | 40219 | 40
2010-02-11 | 40220 | 41
2010-02-12 | 40221 | 42
2010-02-13 | 40222 | 43
2010-02-14 | 40223 | 44
2010-02-15 | 40224 | 45
2010-02-16 | 40225 | 46
2010-02-17 | 40226 | 47
2010-02-18 | 40227 | 48
2010-02-19 | 40228 | 49
2010-02-20 | 40229 | 50
2010-02-21 | 40230 | 51
2010-02-22 | 40231 | 52
2010-02-23 | 40232 | 53
2010-02-24 | 40233 | 54
2010-02-25 | 40234 | 55
2010-02-26 | 40235 | 56
2010-02-27 | 40236 | 57
2010-02-28 | 40237 | 58
2010-03-01 | 40238 | 59
2010-03-02 | 40239 | 60
2010-03-03 | 40240 | 61
2010-03-04 | 40241 | 62
2010-03-05 | 40242 | 63
2010-03-06 | 40243 | 64
2010-03-07 | 40244 | 65
2010-03-08 | 40245 | 66
2010-03-09 | 40246 | 67
2010-03-10 | 40247 | 68
2010-03-11 | 40248 | 69
2010-03-12 | 40249 | 70
2010-03-13 | 40250 | 71
2010-03-14 | 40251 | 72
2010-03-15 | 40252 | 73
2010-03-16 | 40253 | 74
2010-03-17 | 40254 | 75
2010-03-18 | 40255 | 76
2010-03-19 | 40256 | 77
2010-03-20 | 40257 | 78
2010-03-21 | 40258 | 79
2010-03-22 | 40259 | 80
2010-03-23 | 40260 | 81
2010-03-24 | 40261 | 82
2010-03-25 | 40262 | 83
2010-03-26 | 40263 | 84
2010-03-27 | 40264 | 85
2010-03-28 | 40265 | 86
2010-03-29 | 40266 | 87
2010-03-30 | 40267 | 88
2010-03-31 | 40268 | 89
2010-04-01 | 40269 | 90
2010-04-02 | 40270 | 91
2010-04-03 | 40271 | 92
2010-04-04 | 40272 | 93
2010-04-05 | 40273 | 94
2010-04-06 | 40274 | 95
2010-04-07 | 40275 | 96
2010-04-08 | 40276 | 97
2010-04-09 | 40277 | 98
2010-04-10 | 40278 | 99
2010-04-11 | 40279 | 100
2010-04-12 | 40280 | 101
2010-04-13 | 40281 | 102
2010-04-14 | 40282 | 103
2010-04-15 | 40283 | 104
2010-04-16 | 40284 | 105
2010-04-17 | 40285 | 106
2010-04-18 | 40286 | 107
2010-04-19 | 40287 | 108
2010-04-20 | 40288 | 109
2010-04-21 | 40289 | 110
2010-04-22 | 40290 | 111
2010-04-23 | 40291 | 112
2010-04-24 | 40292 | 113
2010-04-25 | 40293 | 114
2010-04-26 | 40294 | 115
2010-04-27 | 40295 | 116
2010-04-28 | 40296 | 117
2010-04-29 | 40297 | 118
2010-04-30 | 40298 | 119
2010-05-01 | 40299 | 120
2010-05-02 | 40300 | 121
2010-05-03 | 40301 | 122
2010-05-04 | 40302 | 123
2010-05-05 | 40303 | 124
2010-05-06 | 40304 | 125
2010-05-07 | 40305 | 126
2010-05-08 | 40306 | 127
2010-05-09 | 40307 | 128
2010-05-10 | 40308 | 129
2010-05-11 | 40309 | 130
2010-05-12 | 40310 | 131
2010-05-13 | 40311 | 132
2010-05-14 | 40312 | 133
2010-05-15 | 40313 | 134
2010-05-16 | 40314 | 135
2010-05-17 | 40315 | 136
2010-05-18 | 40316 | 137
2010-05-19 | 40317 | 138
2010-05-20 | 40318 | 139
2010-05-21 | 40319 | 140
2010-05-22 | 40320 | 141
2010-05-23 | 40321 | 142
2010-05-24 | 40322 | 143
2010-05-25 | 40323 | 144
2010-05-26 | 40324 | 145
2010-05-27 | 40325 | 146
2010-05-28 | 40326 | 147
2010-05-29 | 40327 | 148
2010-05-30 | 40328 | 149
2010-05-31 | 40329 | 150
2010-06-01 | 40330 | 151
2010-06-02 | 40331 | 152
2010-06-03 | 40332 | 153
2010-06-04 | 40333 | 154
2010-06-05 | 40334 | 155
2010-06-06 | 40335 | 156
2010-06-07 | 40336 | 157
2010-06-08 | 40337 | 158
2010-06-09 | 40338 | 159
2010-06-10 | 40339 | 160
2010-06-11 | 40340 | 161
2010-06-12 | 40341 | 162
2010-06-13 | 40342 | 163
2010-06-14 | 40343 | 164
2010-06-15 | 40344 | 165
2010-06-16 | 40345 | 166
2010-06-17 | 40346 | 167
2010-06-18 | 40347 | 168
2010-06-19 | 40348 | 169
2010-06-20 | 40349 | 170
2010-06-21 | 40350 | 171
2010-06-22 | 40351 | 172
2010-06-23 | 40352 | 173
2010-06-24 | 40353 | 174
2010-06-25 | 40354 | 175
2010-06-26 | 40355 | 176
2010-06-27 | 40356 | 177
2010-06-28 | 40357 | 178
2010-06-29 | 40358 | 179
2010-06-30 | 40359 | 180
2010-07-01 | 40360 | 181
2010-07-02 | 40361 | 182
2010-07-03 | 40362 | 183
2010-07-04 | 40363 | 184
2010-07-05 | 40364 | 185
2010-07-06 | 40365 | 186
2010-07-07 | 40366 | 187
2010-07-08 | 40367 | 188
2010-07-09 | 40368 | 189
2010-07-10 | 40369 | 190
2010-07-11 | 40370 | 191
2010-07-12 | 40371 | 192
2010-07-13 | 40372 | 193
2010-07-14 | 40373 | 194
2010-07-15 | 40374 | 195
2010-07-16 | 40375 | 196
2010-07-17 | 40376 | 197
2010-07-18 | 40377 | 198
2010-07-19 | 40378 | 199
2010-07-20 | 40379 | 200
2010-07-21 | 40380 | 201
2010-07-22 | 40381 | 202
2010-07-23 | 40382 | 203
2010-07-24 | 40383 | 204
2010-07-25 | 40384 | 205
2010-07-26 | 40385 | 206
2010-07-27 | 40386 | 207
2010-07-28 | 40387 | 208
2010-07-29 | 40388 | 209
2010-07-30 | 40389 | 210
2010-07-31 | 40390 | 211
2010-08-01 | 40391 | 212
2010-08-02 | 40392 | 213
2010-08-03 | 40393 | 214
2010-08-04 | 40394 | 215
2010-08-05 | 40395 | 216
2010-08-06 | 40396 | 217
2010-08-07 | 40397 | 218
2010-08-08 | 40398 | 219
2010-08-09 | 40399 | 220
2010-08-10 | 40400 | 221
2010-08-11 | 40401 | 222
2010-08-12 | 40402 | 223
2010-08-13 | 40403 | 224
2010-08-14 | 40404 | 225
2010-08-15 | 40405 | 226
2010-08-16 | 40406 | 227
2010-08-17 | 40407 | 228
2010-08-18 | 40408 | 229
2010-08-19 | 40409 | 230
2010-08-20 | 40410 | 231
2010-08-21 | 40411 | 232
2010-08-22 | 40412 | 233
2010-08-23 | 40413 | 234
2010-08-24 | 40414 | 235
2010-08-25 | 40415 | 236
2010-08-26 | 40416 | 237
2010-08-27 | 40417 | 238
2010-08-28 | 40418 | 239
2010-08-29 | 40419 | 240
2010-08-30 | 40420 | 241
2010-08-31 | 40421 | 242
2010-09-01 | 40422 | 243
2010-09-02 | 40423 | 244
2010-09-03 | 40424 | 245
2010-09-04 | 40425 | 246
2010-09-05 | 40426 | 247
2010-09-06 | 40427 | 248
2010-09-07 | 40428 | 249
2010-09-08 | 40429 | 250
2010-09-09 | 40430 | 251
2010-09-10 | 40431 | 252
2010-09-11 | 40432 | 253
2010-09-12 | 40433 | 254
2010-09-13 | 40434 | 255
2010-09-14 | 40435 | 256
2010-09-15 | 40436 | 257
2010-09-16 | 40437 | 258
2010-09-17 | 40438 | 259
2010-09-18 | 40439 | 260
2010-09-19 | 40440 | 261
2010-09-20 | 40441 | 262
2010-09-21 | 40442 | 263
2010-09-22 | 40443 | 264
2010-09-23 | 40444 | 265
2010-09-24 | 40445 | 266
2010-09-25 | 40446 | 267
2010-09-26 | 40447 | 268
2010-09-27 | 40448 | 269
2010-09-28 | 40449 | 270
2010-09-29 | 40450 | 271
2010-09-30 | 40451 | 272
2010-10-01 | 40452 | 273
2010-10-02 | 40453 | 274
2010-10-03 | 40454 | 275
2010-10-04 | 40455 | 276
2010-10-05 | 40456 | 277
2010-10-06 | 40457 | 278
2010-10-07 | 40458 | 279
2010-10-08 | 40459 | 280
2010-10-09 | 40460 | 281
2010-10-10 | 40461 | 282
2010-10-11 | 40462 | 283
2010-10-12 | 40463 | 284
2010-10-13 | 40464 | 285
2010-10-14 | 40465 | 286
2010-10-15 | 40466 | 287
2010-10-16 | 40467 | 288
2010-10-17 | 40468 | 289
2010-10-18 | 40469 | 290
2010-10-19 | 40470 | 291
2010-10-20 | 40471 | 292
2010-10-21 | 40472 | 293
2010-10-22 | 40473 | 294
2010-10-23 | 40474 | 295
2010-10-24 | 40475 | 296
2010-10-25 | 40476 | 297
2010-10-26 | 40477 | 298
2010-10-27 | 40478 | 299
2010-10-28 | 40479 | 300
2010-10-29 | 40480 | 301
2010-10-30 | 40481 | 302
2010-10-31 | 40482 | 303
2010-11-01 | 40483 | 304
2010-11-02 | 40484 | 305
2010-11-03 | 40485 | 306
2010-11-04 | 40486 | 307
2010-11-05 | 40487 | 308
2010-11-06 | 40488 | 309
2010-11-07 | 40489 | 310
2010-11-08 | 40490 | 311
2010-11-09 | 40491 | 312
2010-11-10 | 40492 | 313
2010-11-11 | 40493 | 314
2010-11-12 | 40494 | 315
2010-11-13 | 40495 | 316
2010-11-14 | 40496 | 317
2010-11-15 | 40497 | 318
2010-11-16 | 40498 | 319
2010-11-17 | 40499 | 320
2010-11-18 | 40500 | 321
2010-11-19 | 40501 | 322
2010-11-20 | 40502 | 323
2010-11-21 | 40503 | 324
2010-11-22 | 40504 | 325
2010-11-23 | 40505 | 326
2010-11-24 | 40506 | 327
2010-11-25 | 40507 | 328
2010-11-26 | 40508 | 329
2010-11-27 | 40509 | 330
2010-11-28 | 40510 | 331
2010-11-29 | 40511 | 332
2010-11-30 | 40512 | 333
2010-12-01 | 40513 | 334
2010-12-02 | 40514 | 335
2010-12-03 | 40515 | 336
2010-12-04 | 40516 | 337
2010-12-05 | 40517 | 338
2010-12-06 | 40518 | 339
2010-12-07 | 40519 | 340
2010-12-08 | 40520 | 341
2010-12-09 | 40521 | 342
2010-12-10 | 40522 | 343
2010-12-11 | 40523 | 344
2010-12-12 | 40524 | 345
2010-12-13 | 40525 | 346
2010-12-14 | 40526 | 347
2010-12-15 | 40527 | 348
2010-12-16 | 40528 | 349