In-water measurement of optical backscattering at three angles

J. Ronald V. Zaneveld

Department of Research, WET Labs, Inc.

Philomath, OR  97370

Michael S. Twardowski

Department of Research, WET Labs, Inc.

Narragansett, RI  02874

Casey C. Moore, Wes Stubhar

WET Labs, Inc.

Philomath, OR  97370

(prepared for submission to Applied Optics)

October 5, 2000

Corresponding author:  M.S. Twardowski

Department of Research

WET Labs, Inc.

P.O. Box 468

Saunderstown, RI  02874-0468  USA

phone: (401)783-1787

fax: (401)783-0309

E-mail: mtwardo@wetlabs.com

1.  Introduction

An important parameter in bio-optical algorithms is scattering, particularly scattered light in the backward direction (MODIS algorithms publication).  Water-leaving radiance is composed primarily of backscattered light, with some light directed toward the surface by inelastic processes (e.g., fluorescence).  Measurements of the volume scattering function (VSF) are therefore required to validate remote sensing algorithms and to understand the distribution of upwelled radiation governed by radiative transfer processes in aquatic systems.

The remote sensing reflectance just below the surface, Rrs(0-), is often given as (Morel and Gentili 1993, 1996):

Rrs(0−)=Lu(0−)Ed(0−) = fQ bba.

where Lu is the upwelling radiance, Ed the downwelling irradiance, bb the backscattering coefficient, a the absorption coefficient, and the parameter f/Q is a bidirectional distribution function that is dependent on the shape of the light field and the shape of the volume scattering function (Morel and Gentili 1993;  Zaneveld 1995).  Using a proximate average (such 0.0922, Morel and Gentili 1993) or look-up tables (Morel and Gentili 1996 I think) for f/Q, the dependence of Rrs on the inherent optical properties can be investigated with a knowledge of bb and a only (Morel and Gentili 1993; Morel et al. 1995).

A potentially more accurate estimate of f/Q can be obtained with additional knowledge of the shape of the VSF in the backward direction.  Zaneveld (1995) has shown from radiative transfer theory that f/Q can be approximated by

fQ = β(π−θs)μdbb(1+w) ,

where (s) is the volume scattering function at the angle between the underwater solar zenith angle s and the vertical.The parameterdis the downwelling average cosine of the light field and

w =  1−0.52ω0−0.44ω021−ω0  ,

where 0 is the single scattering albedo (b/c).  It is thus seen that f/Q is directly proportional to the VSF at a particular angle in the backward direction.  By measuring the VSF in the backward with sufficient angular resolution to determine both the scattering at this angle and the full backscattering coefficient, it is expected that closure between the inherent optical properties and reflectance can be investigated with a high level of accuracy.

In addition to remote sensing applications, the magnitude and angular distribution of scattered radiation in the backward hemisphere is valuable in investigations of particle composition in seawater.  Several inversion algorithms have been devised to interpret volume scattering in terms of the composition (bulk refractive index) and size distribution of the particle population.  These are the fundamental parameters needed to describe particle sedimentation and resuspension processes in aquatic systems.

While the need for better understanding the VSF is increasingly acknowledged, measurements characterizing the distribution of scattered light have been until now limited.  In fact, data collected by Petzold (1972) in three representative water masses (San Diego Harbor, Coastal Califiornia, and the Sargasso Sea) still constitutes the authoritative source for in-water VSFs.  Moreover, backscattering coefficient estimates have predominantly been derived from single angle measurements and accordingly have been required to assume a constant relationship between this measurement and bb in all water types (Maffione and Dana 1997).

To address these needs, a sensor, the ECO-VSF, has been developed to estimate the backscattering coefficient by measurement of the volume scattering function at 100, 125, and 150 degrees.  The sensor measures scattering at one wavelength, which can be 440 nm, 532 nm, or 650 nm.  The sensor samples at a rate of 1 Hz, thus providing this information at spatial and temporal scales similar to other instruments engaged in profiling or towed applications.

Herein, we describe the principles of operation and evaluate calibration and measurement protocols for the ECO-VSF.  Field performance is evaluated with measurements made at the Rutgers LEO-15 site off the coast of New Jersey during July of 2000.

2. Sensor description (needs basic work)

3 sources, 1 detector, etc.

<Engineering description here>

Background counts are detected by sensor and are unique to each angle measurement.  These are referred to as dark offsets, Od().  Dark offsets are (. . . need an engineer description of why there is a raw signal without any photon capture).  Dark offsets are measured for each angle by collecting a record while pointing the source beams toward a black velvet cloth about 1 to 2 m away in a completely dark room.

Backscattering instruments, in their essential form, consist of a light source projecting into the water and a receiver placed in proximity to the source, that collects light scattered into an acceptance aperture.  Figure 1 shows the geometry for the ECO-VSF.  The detector (D) and the source (S) are separated by a distance SD.  The angle of the center of the detector beam with the line SD is d.  The half-angle of the detector cone is d.  The corresponding angles for the source are s and s.  S is the origin of the system.

3.  Sensor calibration

The scattering meter measurement may be defined as (ignoring spectral dependencies):

β(θ,Δθ)=2π0πβ(θ)W(θ)sinθdθ,

where θ is the centroid angle computed as:

θ=0πθW(θ)sinθdθ0πW(θ)sinθdθ,

and  is the full-width-half-maximum bandwidth of the weighting function, W(), multiplied by 2sin().  The weighting function describes the probability distribution function for a given scattering measurement based on the geometry of the sensor.

The objective of calibrating a scattering sensor is to obtain a scaling factor, F, that can be used to convert raw counts measured by a detector into β(θ), the volume scattering function evaluated at a centroid angle, θ, for a specific weighting function.  To achieve this, W() must be known and raw counts from scattering must be measured in a solution of known () (e.g., a solution of microspherical beads with known size distribution).  Alternatively, a method to derive W() has been proposed by Maffione and Dana (1997) where a planar Lambertian-reflecting target is passed through the sample volume of the sensor.

Both the above methods for calibration were attempted.  The method involving the microspherical beads (hereafter called ‘the bead method’) is presented in section 3.1.  Weighting functions for each angle were determined numerically based on the geometry of the sensor.  Scattering measurements were then made in a solution of microspherical beads suspended in clean water.  Mie theory was used to calculate ().  The reflecting target method (hereafter called ‘the plaque method’) is presented in section 3.2.

3.1 The bead method

3.1.1 Determination of the weighting function, W()

3.1.1.1 Parameterizaton of the sample volume

The volume above the SD line is broken up into the small volumes xyz  (refer to Fig. 1).  z is in the same plane as the SD line.  From simple geometry we see that:

rs(x,y,z)=x2+y2+z2,rd(x,y,z)=x2+(SD−y)2+z2,cos(π−γ)=rd2+rs2−SD22rdrs.

The scattering angle between the ray from the source to V and the ray from V to the detector is defined by .  Additionally,

tanαs=zx2y2,tanαd=z(SD−y)2+x2.

and

and

The solid angle of the source beam is

 s =   (tan  s)2 ,

and for the detector beam,

 d =   (tan  d)2 .

One important calculation is to discover whether or not V is illuminated by the source beam and whether or not it is within the field of view of the detector. Let us look at the source beam and its projects onto a plane defined by z. The intersection of a plane and a cone is an ellipse (if the illuminated area of the plane is finite). We will determine the equation of the illuminated ellipse at level z (see Fig. 3).

The major and minor axes of the ellipse will be given by amaj and amin.

SFz=cot(θs−σs),SEz=cot(θs−σs),2amaj=z[cot(θs−σs)−cot(θsσs)].

and

The minor axis does NOT occur at the center of the cone, C, but rather at the middle of the ellipse, marked M.  The projection of C on the SD plane is C and the projection of M is M.  Hence,

C S = z cot s ,   and

SM = SE + amaj.

Thus, CM = CM = SM – CS = SE + amaj – z cot s.  Using a side view of the source beam, we can see that:

CRCS=tanσs,CR=CStanσs,CS=zsinθs,CR=ztanσssinθs.

and

The equation of our ellipse is:

x2amin2+yamaj2=1,   with (x,y) = 0 at M.

We already know amaj.  We need to determine amin for x = CR  y = - MC at the ellipse. Substitute these into the ellipse equation to get:

amin2=CR21−[CMamaj]2.

The center of the ellipse is at x = 0  y = SM.  The equation of the ellipse is thus:

x2amin2+(y−SM')2amaj2=1.

All the parameters have already been derived. Any volume V at (x, y, z) is illuminated if the substitution of (x,y) in the left hand side (LHS) gives LHS 1.

Derivation for the detector ellipse is similar.  Fig. 5 shows the geometry.  We will use “d” subscript to indicate detector parameters:

amajd=zZ[cot(θd−σd)−cot(θd+σd),DGd=zcot(θd+σd),DMd'=DGd+amajd,CRd=tanσdcosθd,CMd=DMd'−zcotθd,amin2=CRd21−[CMdamajd]2.

and

and

The equation of the detector ellipse at level z is:

xamind2+[y−(SD−DMd')]2amajd2=1.

All points (x,y,z) within the field of view have a Left Hand Side (LHS) 1.

It is of use in the programming to determine the limits of the illuminated and detected volume. Figure 6 shows the geometry:

ZmaxDT=tan(θd−σd),ZminTS=tan(θs−σs).

and

DT and TS are equal to SD, which is known.  Thus,

Zmin=SD1tan(θd−σd)+1tan(θs−σd),Zmax=SD1tan(θd+σd)+1tan(θs+σs).

The source maximum, Ymaxs, is equal to SEs + 2 amajs; the source minimum, Ymins, is equal to SEs; the detector maximum, Ymaxd, is equal to SG and (SD – DGd); the detector minimum, Ymind, is equal to SH and [(SD – DGd) – 2a majc].  The maximum of x is the smaller of amind or amins.

# The volume of interest is thus defined by,

{										 }

# x: 0 to maximum x (we will multiply by 2 to get the complete volume)

# y: the larger of min Ys and min Yd to the smaller of max Ys and max Yd

z: Zmin to Zmax.

3.1.1.2 Signal strength from an elementary volume, V

Let the source power be Ps.  Ps is spread over a solid angle s =  tan2 s.  What is the irradiance at V?

An area V = x y at an angle 90- s (see Fig. 2) to the center of the source beam has a solid angle of

Ω(ΔV)=ΔxΔycos(90−α(x,y,z))rs2(x,y,z).

The power received by V (x,y,z) is

P(ΔV)=PsΩ(ΔV)Ωs.

The irradiance, or power per unit area is P(V) / xy. This power is attenuated by e-crs(x,y,z).  Thus, the irradiance at (x, y, z) is

E(x,y,z)=P(ΔV)ΔxΔy=Pscos(90−α(x,y,z))rs2(x,y,z)Ωse−crs(x,y,z).

The scattering angle between s, v (x, y, z), and D is  (x, y, z).

Since, by definition, β(γ)=dI(γ)EdV,dI(γ)=β(γ)EdV.by rearrangement

Hence dI ( (x, y, z)) =  ( (x, y, z)) E (x,y,z) xyz, units of Watts/Steradian.

dI = P/d (x,y,z).

The power received at the detector due to scattering at V is then:

Pd (x, y, z) = dI ((x, y, z)) d (x, y, z) e-crd (x, y, z).

d is the detector solid angle for V:

Ωd(x,y,z)=ΔxΔycos(90−αd(x,y,z))rd2(x,y,z).

3.1.1.3 Solving for W()

By substitution, we obtain:

and

Pd(x,y,z)=β(γ(x,y,z))E(x,y,z)ΔxΔyΔz×ΔxΔycos(90−αd(x,y,z))rd2(x,y,z)e−crd(x,y,z),

Pd(x,y,z)Ps=β(γ(x,y,z))W(x,y,z),

where W(x, y, z) is the weighting function for V (x, y, z) with the associated scattering angle .

We now need to sum up all W(x, y, z) for angles  to get the total weighting function. We will normalize this to the total sum of all weights, so that

xyzW(x,y,z)=1.

Calculations to determine the weighting function were made in MATLAB.  The resulting weighting functions for the ECO-VSF sensor are plotted in Fig. 7.

3.1.2 Determining theoretical phase functions

Once the weighting function has been determined, a hydrosol with a known volume scattering function must be used to calibrate the sensor.  Such a hydrosol is microspheres (polystyrene beads) of known size distribution and refractive index suspended in clean water.  Phase functions, β(θ) (=()/b), for such populations of spherical, homogeneous particles are determined with Mie Theory.  By convolving the computed phase function with the weighting functions (section 3.1), the appropriately weighted phase function evaluated for each measurement angle, β(θi) (for i = 100, 125, and 150 degrees), can be determined.  Microspheres have been similarly used in calibrating other scattering devices (e.g., Volten et al. 1998; Voss et al. AO; W***** et al. 1998 L&O).

Mie Theory coded from Bohren and Huffman (1983) and translated into MATLAB (Emmanuel Boss, Oregon State University) was used for the phase function determinations.  Wavelength-dependent refractive indices relative to air for polystyrene were computed from:

np = 1.5663 + 0.00785-2 + 0.000334-4,     with  in m.

Particle size probability distribution functions (area=1) were modeled according to a Gaussian using the mean size and standard deviations reported with the beads from the manufacturer (Duke Scientific).  Reported size distributions were verified with some bead samples using a Coulter Counter device.  Phase functions were computed at typically 500 evenly spaced particle sizes spanning ±3 standard deviations from the mean.  The phase function for each particle size was then weighted according to the distribution function, and summed to give the phase function for the entire population.

3.1.3 Experimental calibration and application

Beads of approximately 2 m with a relatively narrow distribution (±1-2%) were primarily used in calibrations.  Direct measurements with an ECO-VSF were made of raw counts of scattering, N(i,C) for i = 100, 125, and 150 degrees, for the calibration bead population at several concentrations, C, with concurrent measurements of total scattering, b(C) (from a WET Labs ac-9).  The raw scattering counts (ECO-VSF) and total scattering (ac-9) of clean water (before the addition of any beads) were subtracted from the same measurements with the beads to obtain the scattering that was associated with the beads only.  The dark offset, Od, (see section 2) was included in this subtraction.  For each concentration of beads, the slope between the raw counts for each angle and total scattering, [ N(i,C)/b(C) ], was used as the experimental phase function, βe(θi), in terms of counts (m-1)-1.  Note that the concentrations of the beads do not need to be known in this calibration method.

Calibration scaling factors, F(i) were then obtained by dividing the theoretically derived β(θi) (see section 3.2) by the experimental βe(θi).  At this point, basic calibration of the sensor is complete.

For an unknown solution, raw counts from the sensor are converted to a volume scattering function appropriate weighted for each measurement angle, β(θi) of units m-1 sr-1 as follows:

- Subtract the dark offset, Od, and

- Multiply the resulting values by the calibration scaling factor, F(i).

Note that corrected measurements of β(θi) include pure water.  Thus, the full relationship between the raw counts (after dark offset subtraction) and β(θi) is:

N(i) F(i)  =  N(i) β(θi)βe(θi)  =  N(i) β(θi)N(i) / b]

=  b β(θi)β(θi)

3.2 The plaque method

The details of the plaque method have been reported previously in Maffione and Dana (1997).

4. Dependence of the scattering signal on the beam attenuation coefficient

There is attenuation of the incident beam from the light source to the sample volume and the light that is scattered in the sample volume towards the detector.  This results in our scattering measurements being underestimates of true scattering in the water.

Each elementary scattering volume has a specific pathlength from the light source to the detector.  This pathlength and the beam attenuation coefficient (c, m-1) determine the effect of attenuation on the light received from that elementary volume.  Integrating over all elementary volumes gives the total dependence on the bean attenuation coefficient.  In the case of the ECO-VSF sensors, the pathlength is very short, so that c- dependence is significant for large values of c only.  Figure 8 shows this dependence.

(This assumes single scattering)

The dependence on c can readily be modeled, so that the measured scattering function at a given value of c, m(, c), can be corrected to the value for c = 0 m-1 as follows:

m(100, c) e0.0314c = m(100, c = 0);

m(125, c) e0.0441c = m(125, c = 0);

m(150, c) e0.0804c = m(150, c = 0);

Calculations show that even for c = 10 m-1, the centroid angles shift by much less than 1 degree, so that this effect can effectively be ignored.

5. Derivation of total backscattering from volume scattering evaluated at 3 angles

Backscattering coefficients, bb, were obtained in two steps:

- Convert β(θi) measurements to a polar area, A(θi), by multiplying by 2πsin(θi).  Because 2sin() = 0, four values of A are actually known in the backward hemisphere, i at 100, 125, 150, and 180 degrees, where A(180) = 0.

- Fit a third order polynomial to A(θi), and integrate from /2 to .

Testing this approach with volume scattering functions measured by Petzold (1972), errors in bb estimates using only the three centroid angles of the ECO-VSF were minimized by multiplying by 1.013.  Resulting maximum errors for all the Petzold functions, excluding filtered freshwater and filtered seawater, were 1.1% (Fig. 9).

6. Performance in the field (needs basic work)

-Can show that bb/b is reasonable and agrees with theory (Twardowski et al. submitted).

-Can discuss coherence issues between the different measurements – temporally and spatially decoupled; coherence levels at about 0.5 m.

-Further validation with a bench top instrument (e.g., Wyatt Dawn) would be nice, but no data in hand.

7. Discussion and applications

8. References

9. Acknowledgments

Figure 1

D, detector

S, source

d

s

d

s

z

y

x

V =xyz

z

Figure 2



W

W()

9.81.6 m

bead

2.00.022 m

bead

W100()/3

W125()/3

W150()/3

Figure X

Figure  8.  The dependence of the integrated weighting volume on the beam attenuation coefficient.  This dependence is almost exactly log-linear and can readily be modeled (see text). The upper curve is for the 100 sensor, the middle curve for the 125 sensor and the bottom curve for the 150 sensor.

Figure 9.

ECO-VSF data

Rutgers LEO-15 site off the New Jersey coast, USA

temperature (degC)

July 24, 2000

ECO-VSF data

Rutgers LEO-15 site off the New Jersey coast, USA

temperature (degC)

July 26, 2000 (Intercalibration site, NODE B)

ECO-VSF data

Rutgers LEO-15 site off the New Jersey coast, USA

temperature (degC)

July 28, 2000