Prepared by:	Satlantic Incorporated

3481 North Marginal Road,

Richmond Terminal, Pier 9

Halifax, Nova Scotia

B3K 5X8

Tel (902)492-4780

Fax (902)492-4781

ProSoft User Manual 7.7

Document: SAT-DN-00228, Rev. B

# Table of Contents

Table of Contents	ii

List of Tables	v

List of Equations	vi

1.	Introduction	1

2.	System Requirements	2

3.	Installation	3

4.	Revisions	4

4.1	ProSoft 7.7	4

4.2	ProSoft 7.6	4

4.3	ProSoft 7.5	5

4.4	ProSoft 7.4	5

4.5	ProSoft 7.3	5

Overview	7

Processing Context	7

5.1.1	Current Instrument	8

5.1.1.1	New	8

5.1.1.2	Import	8

5.1.1.3	Edit	8

5.1.1.4	Delete	8

5.1.2	Current Parameters	8

5.1.2.1	New	8

5.1.2.2	Edit	8

5.1.2.3	Delete	8

5.2	Multi-Level Processing	9

5.3	Single Level Processing	9

5.4	Tools Menu	9

5.4.1	Ascii Data Extractor	9

5.4.2	HDF Viewer	9

5.5	Help Menu	9

5.5.1	Definitions	9

5.5.2	About	9

5.6	File Menu	9

5.6.1	Exit	9

Instrument Context	10

6.1	Available Calibration Files	10

6.2	Loaded Calibration Files	11

6.3	Calibration File Parameters	11

6.3.1	Sensors	11

6.3.2	Frame Tag	11

6.3.3	Instrument Type	11

6.3.4	Immersion Coefficient	11

6.3.5	Measurement Mode	11

6.3.6	Frame Type	11

6.4	Sensor Parameters	12

6.4.1	Distance to Surface	12

6.4.2	Distance to Pressure	12

6.4.3	Channels	12

6.5	Configuring Sensor Distances	13

6.6	Creating a New Instrument Context	15

Parameters Context	16

8.	ASCII Data Extraction	19

HDF Data Viewer	20

9.1	File Menu	21

9.1.1	Open	21

9.1.2	Save As	21

9.1.3	Print	21

9.1.4	Exit	21

9.2	Attributes Menu	21

9.2.1	HDF File	21

9.2.2	Sensor Group	21

9.2.3	Sensor Data Table	21

9.3	HDF File Selected	22

9.4	Sensor Group	22

9.5	Independent/Dependant Variables	22

9.6	Graph Type	22

9.7	Graphing Options	22

9.7.1	Overlay	22

9.7.2	Grid	22

9.7.3	Rotate	22

9.7.4	Zoom	22

9.7.5	Graph	23

9.7.6	Legend	23

10.	Data Processing Equations	24

10.1	Level 1a - Level 1b Data Processing	25

10.1.1	Application of Calibration Data to Level 1a Files	25

10.1.1.1	Optical Data Calibration	25

10.1.1.2	CAL darks	25

10.1.1.3	LOG darks	25

10.1.1.4	NUL darks	26

10.1.1.5	BIN darks	26

10.1.1.6	Dark Current Correction of hyperspectral (OPTIC3) Data	26

10.2	Level 1b – Level 2 Data Processing	27

10.2.1	Dark Data Deglitching	27

10.2.2	Data Deglitching	28

10.2.3	Profiler Data Level 1b - Level 2 Processing	28

10.2.3.1	Synchronization of optical data frames with pressure at Hyperspectral Profiler	28

10.2.3.2	Pressure TARE Correction	28

10.2.3.3	Tilt Edit	28

10.2.3.4	Reference Instrument	29

10.3	Level 2 – Level 3a Processing	29

10.3.1	Read Level 2 Data	30

10.3.2	Calculate Master Co-ordinates	30

10.3.2.1	Calculate Pressure Co-ordinates	30

10.3.2.2	Calculate Time Co-ordinates	31

10.3.3	Coordinate Interpolation	31

10.3.4	Write Level 2s HDF File	32

10.3.5	Wavelength Interpolation	32

10.3.6	Natural Log Transform	32

10.3.7	Average Data	32

10.4	Level 4 Data Processing	33

10.4.1	Diffuse Attenuation Coefficient	33

10.4.1.1	Integration Points	34

10.4.2	Propagate Optical Variables to Surface	36

10.4.2.1	Reflection Albedo	37

10.4.3	Water Leaving Radiance	37

10.4.3.1	Reflectance Index	37

10.4.3.2	Refractive Index	38

10.4.4	Surface Remote Sensing Reflectance's	38

10.4.5	Remote Sensing Reflectance Profile	38

10.4.6	Surface Reflectance's	39

10.4.7	Reflectance Profile	40

10.4.8	Photosynthetically Available Radiation	40

10.4.9	Chlorophyll a Profile Estimates Morel 98 Model	41

10.4.10	Chlorophyll a Surface Estimates SeaBAM OC2 Model	41

10.4.11	Chlorophyll a Surface Estimates Gordon 88 Model	42

10.4.12	Estimation of Energy Fluxes	42

11.	Appendix A: Term Definitions	43

11.1	File Naming Convention	43

11.2	Optical Sensors	43

11.3	Ancillary Sensors	44

12.	Appendix B: Data Formats	45

13.	Appendix C: Profiler/Reference Data Deglitching Function	49

13.1	Background	49

13.2	Description of data distribution with depth	49

13.3	Problem	49

13.4	Approach for the Problem Solution	49

13.5	Examples	50

53

# List of Tables

Table 1 Instrument Context Parameters	12

Table 2 Dark Correction Parameters	17

Table 3 Tilt Editing Parameters	17

Table 4 Data Filtering	17

Table 5 Level 3a Averaging	18

Table 6 Level 4 Constants	18

Table 7 Processing Levels Supported	24

Table 8 File Header	45

Table 9 Level 1-2 Table Descriptors	46

Table 10 GPS Table Descriptors	47

Table 11 Level 4 Table Descriptors	48

# List of Equations

Equation 1 General Calibration Equation	25

Equation 2 NUL Dark	26

Equation 3 BIN Dark	26

Equation 4 Hyperspectral Dark	26

Equation 5 Hyperspectral Data Calibration	26

Equation 6 Hyperspectral Dark Correction	27

Equation 7 Dark Data Deglitching	27

Equation 8 Standard Deviation	27

Equation 9 Pressure Filtering	30

Equation 10 Diffuse Attenuation Coefficient (Ocean Optics Protocols for SeaWiFS eqn. 26 pg. 49)	33

Equation 11 Ocean Optics Protocols for SeaWiFS eqn. 31 pg. 50	33

Equation 12 Austin Petzold 490nm	34

Equation 13 Austin Petzold 520nm	35

Equation 14 Morel 1988	35

Equation 15 Morel 1988 Equation 9	35

Equation 16 Surface Variables (Ocean Optics Protocols for SeaWiFS eqn. 31 pg. 50)	36

Equation 17 Es(0-)	36

Equation 18 Ls(0-)	36

Equation 19 Normalized Water Leaving Radiance (Ocean Optics Protocols eqn. 63 pg. 54)	37

Equation 20 Water Leaving Radiance	37

Equation 21 Ed(0+)	37

Equation 22 Surface Remote Sensing Reflectance	38

Equation 23 Above Surface Ed(0+)	38

Equation 24 Water Leaving Radiance	38

Equation 25 Remote Sensing Reflectance Profile	38

Equation 26 Surface Reflectance	39

Equation 27 Above Surface Ed(0+)	39

Equation 28 Above Surface Eu(0+)	39

Equation 29 Reflectance Profile	40

Equation 30 Photosynthetically Available Radiation	40

Equation 31 Percentage PAR	40

Equation 32 Reference Par	40

Equation 33 Morel 98 Chlorophyll Model	41

Equation 34 Chlorophyll a Surface Estimate SeaBAM OC2 Model	41

Equation 35 Calculation of R Coefficient	41

Equation 36 Chlorophyll a Surface Estimate Gordon 88 Model	42

Equation 37 Chlorophyll a Surface Estimate Gordon 88 Model	42

Equation 38 Estimation of Energy Fluxes	42

# Introduction

ProSoft is a data analysis package for processing optical data collected from oceanographic measurement systems.  The primary goal of the program was to create a package that would be able to process optical data in an automated manner so that processing of data was not subjective and any two investigators would get the same answers from the same data set.  With the increasing number of different optical instrument types that can be used from the autonomous buoys, from ships, or from airborne platforms, a demand for a generalized approach for processing all optical data has emerged. To meet this demand, ProSoft has been created. The most important changes have been related to data and metadata organization, which is now based on the Hierarchical Data Format (HDF 4), developed by The National Center for Supercomputing Applications at University of Illinois at Urbana-Champaign http://hdf.ncsa.uiuc.edu/. At the same time, the principles of data processing have not changed. Any principle changes that will be introduced in the future will be stated explicitly (they can be found in the program under menu ‘Help->About’). ProSoft allows users to read, calibrate, average and inspect data sets collected from their instrumentation.

The program is currently supported by Satlantic Inc. and is available at ftp://ftp.satlantic.com/pub/sensors/software/ProSoft/

# System Requirements

ProSoft's source code is completely written in MatLab® version 6.5.1. Thus the system requirements for ProSoft are defined by the system requirements of MatLab, which are as follows:

OS: Microsoft® Windows® 95, Windows 98 (original and Second Edition), Windows Millennium Edition, Windows NT 4.0, Windows 2000 or Windows XP.

Processor: Pentium®, Pentium Pro, Pentium II, Pentium III, Pentium IV or AMD® Athlon based personal computer.

Current version of ProSoft has been compiled on PC with Pentium III processor, under Windows NT.

We suggest using a minimum of 256 MB of RAM.

# Installation

ProSoft is available as a standard Microsoft® Windows® 95/98/NT/XP install and is supplied on CD or as a single self-extracting program from our Internet site at the above ftp address. One should run the ProSoft#.#_Setup.exe program and follow the instructions on the screen.

# Revisions

## ProSoft 7.7

- Changed graphics renderer mode to zbuffer which renders 3-D graphics in much less time.

- When graphing hdf file data, users are now able to select which independent variables to graph as well as the range of the dependant variable.

- New graphical overlay option allows users to graph data from different dependant variables belonging to the same instrument group onto the same graph.

- Any number of ECO Sensor IOP sensors can now be processed for an instrument. Most calibration file definitions are acceptable for an ECO Sensor IOP.

- Improved reliability of sensor distances from surface and pressure reference distances.

- Sample delay time correction has now been applied to timer values where appropriate at level 2 processing.

- Transmissometer sensor has been added to level 2s and 3a data processing.

- Added Reference Ef Ev and Ld optical sensors to level 2s and 3a data processing.

- Any kind of profiler can now be used in Reference mode. The mode is indicated in the instrument context by selecting Reference as the instrument type.

- Improved dynamic data processing at levels 2s and 3a.

- All GPS telemetry definitions are now supported.

## ProSoft 7.6

- GPS sensor data integrated into profiler/reference configurations. GPRMC, GPGLL and GPGGA only are supported.

- TSRB mode added for all profiler/reference configurations including HyperPro II.

- Level 4 water properties data product enabled. Instrument must have a profiler with Temp and Cond sensors.

- Level 2s depth integration resolution is now adjustable to 0.01, 0.02, 0.05 and 0.10m through parameters settings.

- Sas integration time interval is now derived on the interval of the optical sensor with the highest rate instead of the standard of 0.1sec.

- Conductivity sensor data included to level 3a.

- Temperature sensor data included to level 3a.

- Fluorometer sensor data included to level 3a.

- File batch processing sequence changed. Instead of processing all files together at one level (i.e. process all files at level 1a before processing all files at level 1b), each file is processed separately from level 1a to selected level.

- Calibration files can now be added or removed from configuration files that are created from *.sip files.

- Imported configuration files can now be saved with their original file name by clicking on 'Save' in the configuration utility.

- Reference only or TSRB mode data are now integrated onto the time interval derived from the Es sensor instead of the previous standard of 0.1 sec.

- Added Reference Ev optical sensor to processing at all levels.

## ProSoft 7.5

- Added support for level 2s GPS integration for the following GPS formats $GPRMC, $GPGLL and $GPGGA.

- SAS GPS data is now averaged at level 3a.

- GPS data is not available for viewing below level 2s in the hdf viewer but can be extracted using the ascii data extractor.

- Added support for the ECO Series IOP sensor for the new Hyperpro II instrument.

- Faster ascii data extraction of hdf files.

- Fixed bugs preventing processing of TSRB data.

## ProSoft 7.4

- Processing parameters have been incorporated into one file. Access to the file is through the new processing parameters utility.

- Processing parameters can be easily edited through the new parameters interface. All four levels of parameters are viewed and edited at once.

- Processing parameters have been separated from the instrument context. The new ‘Processing Context’ consists of both the current instrument and current processing parameters.

- HDF viewer now includes the ability to save the graph image as a file using the png graphic format.

- Updated ProSoft main menu. Processing level commands have been moved to the main menu for easier use. Ascii data extractor and HDF viewer are now accessible through the menu selection ‘Tools’.

- Instrument context creation/edit utility now allows adding or removing calibration files from the instrument context.

## ProSoft 7.3

- New level 2s file is introduced for all instruments. The Level 2s file shows the interpolated data just prior to averaging at Level 3a.

- Introduction of ‘Instrument Context’ creation and loading for easy data processing and selection between different instruments.

- Added processing support for Satlantic Satnet instruments.

- Easier to use Ascii data extraction utility.

- Updated ProSoft main menu interface.

- Addition of completely revised HDF graphical viewer utility which makes it possible to graph all HDF files in 2-D or 3-D graphical views.

- Updated and easier to use configuration file creation/edit utility.

- Level 4 chlorophyll a profile estimates using Morel 98 model added.

- Level 4 chlorophyll a surface estimates using SeaBAM OC2 model (Rrs) added.

- Level 4 chlorophyll a surface estimates using Gordon 88 model (Lwn) added.

- Level 4 energy profile/surface fluxes added.

- Sas data can now be processed to level 3a.

Loaded parameters

Delete instrument context

Create new instrument context

Create new parameters

Import existing configuration file

Edit current instrument context

Edit current parameters

Process data from raw to level 2s

Process data from raw to level 3a

Process data from raw to level 4

Process data from raw to level 1a

Process data from level 1a to level 1b

Process data from level 1b to level 2s

Process data from level 2s to level 3a

Process data from level 3a to level 4

Delete parameters context

Figure 1 ProSoft Main Menu

Overview

Processing Context

The processing context defines all the parameters necessary for processing of optical data. The processing context is made up of two parts, the current instrument and the current parameters. The current instrument defines the instrument used for gathering the raw data. See instrument context for details The current parameters defines all the processing variables for level 1 to level 4 data processing. See parameters context for details.

### Current Instrument

Displays the currently loaded instrument context. All instrument configuration information is loaded and ready for processing. See instrument context for details.

#### New

Starts the process for creating a new instrument context. The user is asked to supply the location of the calibration files for the instrument, which are then used to create the instrument context. Note: All calibration files for the instrument context must remain in the same directory.

#### Import

Allows the user to import an existing configuration file for use in a new instrument context.

#### Edit

This option loads the current instrument context into the instrument context utility for editing. The user can choose to save the file under the existing instrument context name (‘Save’ ) or create a new instrument context (‘Save As’).

#### Delete

Permanently deletes the current instrument context. The original configuration file associated with the instrument context is not deleted.

### Current Parameters

Displays the currently loaded parameters context. All processing parameters are loaded and ready for processing. See parameters context for details.

#### New

Starts the process for creating a new processing parameters context. The processing parameters utility will display with default values automatically loaded.

#### Edit

This option loads the current processing parameters into the processing parameters utility for editing. The user can choose to save the file under the existing processing parameters context name (‘Save’) or create a new parameters context (‘Save As’).

#### Delete

Permanently deletes the loaded processing parameters context. The original parameters file associated with the processing parameters context is also deleted.

## Multi-Level Processing

Multi-level processing enables the user to process raw data up to level 2s, level 3a or level 4. All intermediate processing levels are automatically processed with accompanying files being produced. To use this feature select the command button according to the ending processing level as desired. ProSoft will prompt the user to select the directory and choose files of raw data.

## Single Level Processing

Single level processing enables the user to process data over a single level. For example from level 1a to level 1b. To use this feature select the command button according to the starting processing level and ending processing level as desired. ProSoft will prompt the user to select the directory and choose files for the starting processing level.

## Tools Menu

### Ascii Data Extractor

This utility converts hdf files to tab separated ascii file format. ProSoft will prompt the user to select the directory and choose files for ascii data extraction. Hdf files at any level except raw (Level 1) can be extracted. See ascii data extraction for details.

### HDF Viewer

Select this option to access the HDF viewer utility which allows the user to view hdf files in 2-D or 3-D graphical format. The utility will open in the empty state. Simply select ‘File -> Open’ to load hdf files for viewing. See hdf data viewer for details.

## Help Menu

### Definitions

Select this option to view a list of definitions of ProSoft terms. See appendix A term definitions for list of definitions.

### About

Select this option to view the history of revisions for ProSoft versions. This lists the latest changes made to ProSoft. See revisions for details.

## File Menu

### Exit

Select this option to exit the current ProSoft session. The last loaded instrument context and processing parameters context will be automatically loaded next time ProSoft is started.

Figure 2 Instrument Context Utility

Instrument Context

Near the top of the main menu is a section labeled ‘Current Instrument’. Instrument Context is defined as the current instrument loaded into ProSoft that will be used for data processing. (The terms instrument context and instrument will be used interchangeably throughout this manual). The instrument context file used to describe this instrument is loaded into memory ready to be recalled when necessary for processing.

The user can switch between instruments easily by clicking on the drop down box and selecting from the list of instruments contexts that have been defined by the user. When an instrument is chosen the instrument context file associated with the instrument is automatically loaded and ProSoft is ready to process data.

## Available Calibration Files

This is a list of all calibration files (*.cal and *.tdf) located in the current directory. In the case of creating a new instrument context the current directory is selected prior to displaying the instrument context utility. When editing or importing the current directory is the directory that the instrument context file is located in.

## Loaded Calibration Files

This is a list of all calibration files (*.cal and *.tdf) that are loaded in the instrument context. These are the calibration files which will be associated with the instrument context.

## Calibration File Parameters

The ‘calibration file parameters’ displays all current configuration information on the calibration file selected in the ‘loaded calibration files’ selection box.

### Sensors

This is a list of all sensors defined within the selected calibration file.

### Frame Tag

Indicates the frame tag associated with the calibration file as it appears in the raw or level 1 telemetry data.

### Instrument Type

Select the type of instrument represented by the calibration file. For a list of possible values see instrument context parameters.

### Immersion Coefficient

Select which calibration coefficients to use for processing depending on instrument  immersion mode (i.e. Air or Water).

### Measurement Mode

Select which mode of measurement to use for processing the calibration file and it’s associated raw data.

### Frame Type

Select the frame type for the calibration file for dark current correction calculations.

Table 1 Instrument Context Parameters

| Instrument Type | Measurement Mode | Frame Type |
| --- | --- | --- |
| Profiler | FreeFall SQM Calibration | ShutterLight ShutterDark Anc LightAncCombined |
| Reference | Surface SQM Calibration | ShutterLight ShutterDark Anc LightAncCombined |
| TACCS | Chain SQM Calibration | LightAncCombined |
| SAS | VesselBorne AirBorne SQM Calibration | ShutterLight ShutterDark Anc LightAncCombined |
| GPS | Default | Default |
| ECO Series IOP | FreeFall Surface | Default |
| For all Instrument Types (except GPS and ECO Series IOP) Media Type are Water or Air. For GPS and ECO Series IOP, Media Type is Default. |  |  |

## Sensor Parameters

The ‘sensor parameters’ displays all current configuration information on the sensor selected in the ‘Sensors’ selection box.

### Distance to Surface

This is the physical distance between the sensor head and the top of the instrument for Profilers, and the distance between the sensor head and the surface of the water, if immersed, for References. SAS instruments do not need to specify distances. See configuring sensor distances for details.

### Distance to Pressure

This is used only for profilers to represent the distance between the sensor head and the pressure reference line on the profiler.

### Channels

All the channels available from the selected sensor are displayed here. For optical sensors this will be a list of all sensor wavelengths.

## Configuring Sensor Distances

The sensor head distances are needed in order to compensate for the fact that the pressure sensor is not located at the same position as the sensor heads. In order to calculate the pressure at the sensor head this geometrical difference must be known and included in the instrument context file (all units must be in meters). When the pressure tare is taken during measurements, two possible methods exist:

1. Pressure tare is taken when the Ed sensor is located at the surface of the water. In this, the most common method, the pressure tare then becomes a combination of the atmospheric pressure and the pressure of the water column from the surface to the pressure sensor reference line. The pressure value at Ed is calculated by subtracting the pressure tare value from the measured pressure values. The pressure value at Lu is calculated by subtracting the pressure tare value from the measured pressure values, as calculated for Ed, and then adding the distance between the Lu sensor and the surface reference.

Pressure Ed = Measured Pressure – Pressure Tare

Pressure Lu = Measured Pressure – Pressure Tare + Lu Distance to Surface

2. Pressure tare is taken when the profiler is located on deck. In this case the pressure tare is equal to the atmospheric pressure. The pressure tare value is then added to the distance from Ed to the pressure sensor reference line to give a new pressure tare value to compute Ed and Lu pressure values using the method described in 1 above.

Pressure Tare = Pressure Tare (Atmospheric Pressure) + Distance Ed Pressure Reference

Pressure Ed = Measured Pressure – Pressure Tare

Pressure Lu = Measured Pressure – Pressure Tare + Lu Distance to Surface

Note: Instrument configurations and distances may not be the same for your instrument. Always measure distances when creating an instrument context file for your instrument. The diagrams below indicate the most common

Atmospheric Pressure

Ed Distance to pressure reference

Pressure Tare

Ed Sensor

Lu Sensor

Lu Distance to surface reference

Pressure reference line

Sea Level

Pressure Reference Line

Ed Distance

Reference: Pressure

0.663575 m

0.101346 m

Lu Distance

Reference: Surface

Figure 3 Profiler Sensor Dimensions

configurations available.

## Creating a New Instrument Context

When using ProSoft for the first time a warning message is given stating that no instrument context has been defined. To create a new instrument context use the following steps:

- Click on ‘New’ in the ProSoft main menu.

- ProSoft will ask the user to point to a directory containing all the calibration files. Note: The calibration files and instrument context file *.cfs should always remain together in the same directory.

- The instrument context utility will then display showing a list of all available calibration (*.cal) and tdf files (*.tdf) located in the directory selected in step 2.

- Highlight all the calibration files in ‘Available Calibration Files’ needed for the instrument context and click on add ‘>>’ to load the calibration files.

- Configure all calibration file and sensor parameters as necessary and click on ‘Save As’.

- The user is then prompted to enter an instrument context name. Once entered click on ‘Ok’.

- A dialogue box will then display confirming that the instrument context has been successfully created. Click on ‘Ok’ and ProSoft will load the instrument context just created.

Whenever the user exits the program, ProSoft remembers the last instrument context that was loaded and automatically reloads that context when starting ProSoft.

Figure 4 Processing Parameters Utility

Parameters Context

Near the top of the main menu is a section labeled ‘Current Parameters’. Parameters context is defined as the parameters that are loaded into memory to be used for data processing. For example they determine what type of dark correction to apply or how many data points to use when averaging.

The user can easily switch between parameters contexts by clicking on the drop down box and selecting from the list of parameters contexts that have been defined by the user. When a parameters context is chosen the parameters file associated with the parameters context is automatically loaded and ready to process data.

A detailed listing of all the processing parameters are given in the tables below.

Table 2 Dark Correction Parameters

| Parameter Name | Parameter Values | Comments |
| --- | --- | --- |
| Auto Dark Correction | CAL BIN LOG SHUTTER | Darks from cal file Darks from profile Darks from data logged with caps on. Darks from shutter frames |
| Dark Bins | Positive Integer | Number of data points that are used to obtain bin darks from profile. Thus Dark Bins is used only if Auto Dark Correction is set to BIN. |
| Shutter Dark Deglitch | ON OFF | ON – First difference filter will be supplied to deglitch shutter dark frames of optical data. |

Table 3 Tilt Editing Parameters

| Parameter Name | Parameter Values | Comments |
| --- | --- | --- |
| Auto Edit | ON OFF | Automated profiler data editing using tilts or falling velocity. ON – automatic editing OFF – interactive editing |
| Tilt Edit | ON OFF | Toggles tilt edit mode on or off. |
| Low Velocity | Positive integer | If tilt edit and auto edit are turned ON, low velocity will be used for automatic editing only if tilt sensor is missing. |
| High Tilt | Positive integer | If tilt edit and auto edit are turned ON, high tilt will be used for automatic tilt editing. |

Table 4 Data Filtering

| Parameter Name | Parameter Values | Comments |
| --- | --- | --- |
| Deglitch Profiler Data | ON OFF | If ON, pofiler data is deglitched using the value of Profiler Noise Threshold. |
| Profiler Noise Threshold | Positive Integer | Adjusts the sensitivity of profiler data deglitching. |
| Upper Depth Level | Positive Integer | Sets shallow depth level below which profiler data is deglitched. |
| Lower Depth Level | Positive Integer | Sets deep depth value above which profiler data is deglitched. |
| Deglitch Reference Data | ON OFF | If ON, reference data is deglitched using the value of Reference Noise Threshold. |
| Reference Noise Threshold | Positive Integer | Adjusts the sensitivity of reference data deglitching. |

Table 5 Level 3a Averaging

| Parameter Name | Parameter Value | Comments |
| --- | --- | --- |
| Bin Interval | Any 0.1m interval (i.e. 0.2, 0.6, 1.2, 2.0 etc) | Controls the depth interval which is the center point for profiler averaging. |
| Bin Width | Any 0.1m interval (i.e. 0.2, 0.6, 1.2, 2.0 etc) | Controls the number of data points used in averaging based on depth. |
| Time Interval | Any 1 second interval (i.e. 1, 2, 10, 17 etc.) | Controls the time interval which is the center point for reference/sas averaging. |
| Time Width | Any 1 second interval (i.e. 1, 2, 10, 17 etc.) | Controls the number of data points used in averaging based on time. |
| Depth Resolution | Variable | Controls the depth resolution used for data integration at level 2s. |
| Wavelength Interpolation | OFF, 1, 2, 5, 10 | Controls the interpolation interval when interpolating onto a constant wavelength. Only applicable for hyperspectral instruments. |

Table 6 Level 4 Constants

| Parameter Name | Parameter Value | Comments |
| --- | --- | --- |
| Integration Points | Must be an odd number (i.e. 1, 3, 5, 7 etc.) | Number of data points used to calculate K values. |
| Reflection Albedo | Default value of 0.043 | Fresnel reflection albedo for irradiance from sun and sky. |
| Reflectance Index | Default value of 0.021 | Fresnel reflectance index of seawater. |
| Refractive Index | Default value of 1.345 | Fresnel refractive index of seawater. |

# ASCII Data Extraction

The ascii data extractor utility allows the user to convert any ProSoft hdf file into tab separated ascii file format. This format can be used to import data into an excel spreadsheet or other program that can use ascii tab delimited files.

To extract hdf files use the following procedure:

- From the main menu click on menu ‘Tools -> Ascii Data Extractor’.

- A folder selection dialogue will appear from which the user selects a folder with the hdf files to be extracted.

- Next a list of all hdf files in the directory selected in step 2 will be displayed. The user can select as many hdf files as desired.

- Click on ‘Ok’ to start the file extraction process.

All the extracted ascii files will be place in a directory called ‘Ascii Files’. This directory will be located in the same directory as the hdf files selected in step 2. If the ‘Ascii Files’ directory cannot be created then the ascii files will be located in the same directory as the hdf files selected in steps 2 and 3.

ProSoft will write up to a maximum of 256 columns per row after which the data will wrap to the next section of rows. This ensures the ascii file can be imported to excel without loss of data.

Included in the ascii file is the hdf file header which includes metadata important for data interpretation.

File menu: Access to printing graph, loading new hdf files or quitting

Attributes menu: Access to hdf file attributes, sensor group attributes or sensor data table attributes

Hdf file selected for viewing

Graph legend showing sensor fields

Graph title showing file name, cruise identifier and processing level

Select sensor group for viewing

Select range of dependant variable

Click to graph sensor fields

Graphing options All Fields, Grid, Rotate, Zoom

Select dependant variable

Select graph type, 2-D or 3-D

Select range of independent variable

Select independent variable

Figure 5 HDF Viewer

HDF Data Viewer

The hdf viewer enables the user to view ProSoft data in graphical format. To use the hdf viewer simply apply the following steps:

- To start, from the main menu, select menu ‘Tools -> HDF Viewer’.

- The HDF Viewer utility will open but with no hdf files loaded. To load files select menu ‘File -> Open’.

- A folder selection dialogue box appears from which the user selects the folder containing hdf files for viewing.

- Next a list of all hdf files in the directory selected in step 2 will be displayed. The user can select as many hdf files as desired then click on ‘Ok’.

The hdf viewer controls are outlined as follows:

## File Menu

### Open

Loads a new set of hdf files for viewing.

### Save As

Save the graph to a file using the png graphic format.

### Print

Prints the graph currently displayed in the graph axes.

### Exit

Exits the hdf viewer utility and returns to the ProSoft main menu.

## Attributes Menu

### HDF File

Select to view the hdf file attributes or metadata applicable to the entire hdf file such as cruise id, date, processing level etc.

### Sensor Group

Select to view the attributes or metadata applicable to the selected sensor group such as instrument type, media, measurement mode etc.

### Sensor Data Table

Select to view the attributes or metadata applicable to the selected sensor data table such as units for each sensor field.

## HDF File Selected

This drop down box displays the current hdf file being viewed. All hdf files that were selected for viewing are listed in this drop down box. To switch files simply click on the drop down box and choose an hdf file.

## Sensor Group

This control lists all the sensor groupings for the selected hdf file. Each sensor group corresponds to an instrument calibration file. Common sensor groups are Profiler, Reference, Sas, ECO Series IOP etc.

## Independent/Dependant Variables

These controls list all the independent and dependent variables available for the chosen sensor group.

## Graph Type

Use the graph type buttons to select which type of graph is required. Note: Certain sensor data tables are not available for viewing in 3-D. In this case the option for selecting 3-D will be grayed out and unavailable for selection.

## Graphing Options

### Overlay

When this option is selected any subsequent graphing will overlay each other. This useful for comparing data from different dependant variables.

### Grid

When this option is selected a grid overlay will be drawn on the graph. This option is only available for 2-D graphs as it is always on by default for 3-D graphs.

### Rotate

When this option is selected the user is able to hold down the left mouse button while over the graph and move the mouse to rotate the graph view. This option is only available for 3-D graphs.

### Zoom

When this option is selected the user is able to zoom in or out of the graph. To zoom in left click the mouse button while over the graph and to zoom out right click the mouse button. To select a portion of the graph to zoom, left click and hold the mouse button while over the graph and drag the area selection box. When completed release the mouse button and the graph will zoom on the area selected. This option is only available for 2-D graphs.

### Graph

When all the desired independent/dependant ranges are selected click on ‘Graph’ to draw the graph in the axes. When zooming or rotating a graph simply click on ‘Graph’ to return the graph to it’s original view.

### Legend

In 2-D graphing mode a legend is displayed on the axes showing the field color-coding scheme for the graph. The legend can be moved by left clicking and dragging the legend to a different location on the axes. Make sure the zoom option is turned off when moving the legend. In 3-D graphing mode a color contour legend is displayed outside the axes, relating color with z-axis value for easier viewing.

# Data Processing Equations

Table 7 Processing Levels Supported

|  | Profiler | Reference | SAS | TSRB | TACCS |
| --- | --- | --- | --- | --- | --- |
| 1a | X | X | X | X | X |
| 1b | X | X | X | X | X |
| 2 | X | X | X | X | X |
| 2s | X | X | X | X |  |
| 3a | X | X | X | X |  |
| 4 | X | X |  | X |  |

Note: Includes both multispectral and hyperspectral instruments

This section is intended to give an overview of the main steps of radiometric data L(z, ) processing carried out by ProSoft.  It is assumed that the radiometric data has been collected using an optical instrument with a raw data format that is in compliance with Satlantic Instrument Files Standard (SIFS).

ProSoft processing is segmented into 4 main levels:

Level 1 - is raw binary data file from an instrument.

Level 1a - binary data is extracted from raw data under the control of the instrument (calibration and/or telemetry definition) files. Extracted information is grouped along with its calibration information and is placed into Level 1a HDF files. File nametag is _L1a.

Level 1b - level1b data is calibrated. No data editing is applied. Shutter darks, if present, are not applied. File nametag is _L1b.

Level 2 - includes Level 1b data, which is further modified per request (i.e. depending on settings of processing parameters and on instrument context). File nametag is _L2.

- Shutter dark correction is applied; reference and dark data deglitching is applied.

- Profiler's data is tilt edited.

Level 2s – level 2 data is interpolated onto a common co-ordinates vector, which is either depth (Profiler) or time (Reference only or Sas). File nametag is _L2s.

Level 3a – includes averaging of Level 2s data as defined by the processing parameters. File nametag is _L3a.

Level 4 – includes several products (users choice) produced from the level 3a data. This includes products such as normalized water leaving radiances, reflectance profiles, photosynthetically available radiation etc. File nametag is _L4.

## Level 1a - Level 1b Data Processing

### Application of Calibration Data to Level 1a Files

#### Optical Data Calibration

Standard optical sensor data formats are processed differently based on the capabilities of the various types of acquisition systems.  These are referred to (in SIFS) as OPTIC1 (high-resolution gain switching 24 bit systems), OPTIC2 (standard 12, 16, 24 or 32 bit systems), and OPTIC3 (hyperspectral systems with adaptive integration).  Application of the calibration data to all optical and ancillary sensors is carried out in accordance to the procedures detailed in SIFS for conversion from binary (or ASCII) digital counts into engineering units.  In general, optical data is converted into engineering units in accordance to the calibration equation:

Equation 1 General Calibration Equation

LLightDatz,λ=LCountsLightDatz,λ−LCalDarkDatλ⋅aλ⋅icλ

where a  is a calibration coefficient and ic is an immersion coefficient obtained from a calibration file.  To simplify notation, in the following dependence on wavelength (λ) will be omitted.  LCalDarkDat⋅a⋅ic is the dark current in engineering units that can be obtained from a calibration file or using one of the dark current correction (DCC) methods described below.

ProSoft currently implements DCC other than calibration dark or shutter dark only in OPTIC1 (high resolution 24 bit systems) fitting mode (see SIFS for detailed explanation of the fitting modes).  Dark current can change under changing thermal conditions on these high-resolution systems.  Dark current correction has to be adjusted accordingly.  ProSoft provides the different DCC methods that will be described in the following sections.

It is important to note that DCC other than based on the calibration file (except hyperspectral) can only be used for the measurement frames obtained with gain switch 1 or higher in OPTIC1 fitting mode (in gain switch 0, or low gain, the CAL darks are used).  In OPTIC1 fitting mode, ProSoft first analyses the measured frames according to the gain switches and applies the DCC according to the Dark Current Correction scheme selected by the user.  In OPTIC2 fitting mode there are only two options available, CAL and NUL.

#### CAL darks

DCC method with calibration darks is given by the general calibration equation.  This is the default mode for both OPTIC1 and OPTIC2 data types.

#### LOG darks

Prior to or just after a light profile (profiler measurement with no caps), darks can be measured by taking a profile with the profiler caps on. Both data files are first processed with the NUL dark option (i.e. no darks subtracted). The LOG dark corrected light profile is then obtained as the difference of the original light profile LLightDat(z,) and the LOG dark values. During level 2 processing ProSoft will prompt the user to select a LOG dark file to use.

#### NUL darks

NUL dark is a special mode in which no darks are subtracted during data calibration (note that in the fitting mode OPTIC1, cal darks will be still subtracted for the frames with gain switch 0).

Equation 2 NUL Dark

LLightDatz,λ=LCountsLightDatz,λ⋅aλ⋅icλ

#### BIN darks

If the profiler reaches a depth for which all the optical sensors reach their dark level, then the darks can be computed from the average of a number of samples at the bottom of the profile.  For each wavelength λ, the value of DCC is obtained from a layer where average minimum light values minLLightDatz,λ are observed.

Equation 3 BIN Dark

LLightDatz,λ=LLightDatz,λ−minLLightDatz,λz=zmin,...,zmax

#### Dark Current Correction of hyperspectral (OPTIC3) Data

Usually hyperspectral data is dark corrected with the values obtained from shutter darks to obtain the most accurate correction.  Shutter darks are continuously recorded during the measurements by occulting the input fiber with an optical shutter, typically after every five light samples.  Hyperspectral calibration and subsequent DCC is carried out in the following steps:

- Correct shutter dark counts obtained from a log file by dark offset (obtained as the difference between shutter darks and capped darks).

Equation 4 Hyperspectral Dark

LCountsDarkDat=LCountsDarkDat−LCountsDarkOffset

- Convert data counts into engineering units in accordance to the calibration equations. The calibration equations for optical hyperspectral data is:

Equation 5 Hyperspectral Data Calibration

LLightDat=LCountsLightDat−LCalDarkDat⋅a⋅icit1it2LDarkDat=LCountsDarkDat−LCalDarkDat⋅a⋅icit1it2

where a  is the calibration coefficient, ic is an immersion coefficient, it1 is the integration time during calibration and it2 is the integration time during the measurement.  a, ic and it1 are taken from a calibration file,  it2 is obtained from the same log file as optical data.

- Deglitch dark data using a first difference filter (optional).

- Interpolate shutter darks as a function of measurement time to match the number of dark and light data measurements.

- Correct light data using shutter darks.

Equation 6 Hyperspectral Dark Correction

L=LLightDat−LDarkDat

Note: SAS Instrument

GPS UTC time that was presented at Level 1a in HHMM (hours, minutes) format is recalculated into seconds from the start of the current day. GPS date that was presented at Level 1a in DDMMYY (day, month, year) format is recalculated into days since start of the current year.

Note: Profiler

If one wants to use other than shutter dark (can be applied only to hyperspectral profiler) correction (e.g. NUL) then AUTODARK settings should be changed for all profiler sensor groups respectively i.e. one should change AUTODARK from SHUTTER to NUL in the profiler sensor group that contains Lu as well as in the profiler group that contains Ed. One will get wrong results if one group will have AUTODARK set to SHUTTER, and the other group to NUL. However, one could use any of CAL, BIN, or NUL dark corrections in any of the groups.

## Level 1b – Level 2 Data Processing

All optical data sensors are checked for wavelength sequence. The columns of optical data are organized from left to right as a function of the increasing wavelength.

Optical data can be also deglitched if respective settings in processing parameters are set. As a rule the deglitching is not required.

### Dark Data Deglitching

Optional deglitching module is applied to data to remove obvious glitches in dark data.  First, a first difference operator is applied to the time series of N frames of optical data at each channel:

Equation 7 Dark Data Deglitching

dLλi=Lλi+1−dLλi,i=1,...,N−1

Second, the standard deviation is calculated for each channel:

Equation 8 Standard Deviation

stdLλ=1N−2i=1N−1dLλi−dLλ2

Third, any data which values are higher than 10stdLλ is blanked by a blanking constant NaN.

### Data Deglitching

This optional module can be toggled on/off using the processing parameters Deglitch Profiler Data and Deglitch Reference Data. For details on how it works see profiler/reference deglitching.

### Profiler Data Level 1b - Level 2 Processing

Profiler data has a number of corrections and checks before the level 2 data files are created.  These include pressure tare (offset) correction, wavelength sequence check and a tilt quality control edit.

#### Synchronization of optical data frames with pressure at Hyperspectral Profiler

Optical data frames are synchronized with pressure using timetags that are appended to each frame by SatView. Once the pressure frames are synchronized with the optical frames, TIMER time values are used to take into account the differences in integration times of the different optical data frames (i.e. Lu integration time may exceed the integration time of Ed). As a result interpolated pressure vector is appended to each sensor which contains any of the optical light frames. Note that if integration times, i.e. for Lu and Ed, are different, the appended pressure vector values also will be slightly different.

#### Pressure TARE Correction

Pressure sensors often have a small offset value that will change with atmospheric pressure.  The offset is removed by a pressure tare.  This value is stored in the header records by SatView when the data is collected.  The pressure tare is defined as the pressure value that is measured when a profiler's Ed sensor is located just at the surface.  This value has to be later subtracted from the measured pressure values of a profile.  If the profiler measurements were logged with SatView, the pressure tare value is automatically inserted into the file header.  If no information is obtained, pressure tare value is set to 0.

#### Tilt Edit

Profiler measurements can be edited interactively or automatically removing any tilt contaminated records.  In manual mode, a plot is shown with profiler tilts (if available) and profiler velocity as a function of depth.  With the interactive graphical editor the user defines an upper zmin and lower depth zmax (we assume that z co-ordinate is directed downwards i.e. depth is zmin surface and zmax at the bottom of a profile).  Data above zmin or below zmax will be removed.  A tilt Tiltz threshold Tiltlimit (default is 5 degrees) usually is used for estimation of zmin and zmax.  In a case of missing tilt sensor, the freefall velocity of the profiler can be used (zero or negative velocity is assumed to be high tilt).  Automated editor module determines zmin and zmax in 10% of the upper and lower part of profiler measurements.  The upper depth and the lower depth are defined as the first depth level (z) where the following conditions are met:

Tiltz<Tiltlimitmin(z)≤z<0.1max(z) for top layer of the profilemax(z)−0.1max(z)<z≤max(z) for bottom layer of the profile

If a reference instrument is present, then the matching records in the reference data are also removed.

#### Reference Instrument

Reference data is checked before the level 2 data files are created.  These checks include wavelength sequence check and a spike deglitch.

## Level 2 – Level 3a Processing

Read Level 2 Data

Calculate Master Co-ordinates

Calculate Pressure

Co-ordinates

Calculate Time

Co-ordinates

Wavelength Interpolation

Co-ordinate Interpolation

Natural Log Transform

Average Data

Write Level 2s File

Figure 6 Level 2-> 3a Data Processing Flow

### Read Level 2 Data

Data from the Level 2 hdf file is read and organized depending on the type of instrument (i.e. Satnet Hyperspectral Profiler/Reference).

### Calculate Master Co-ordinates

At level 3a processing, the averaging of data is done within intervals based on a system of co-ordinates. For systems with a profiler the co-ordinates are based on pressure (averaging in depth). For systems with a reference only or SAS the co-ordinates are based on time (averaging in time).

#### Calculate Pressure Co-ordinates

There exists in most instruments a difference in distance between sensors. For example in an SPMR profiler the distance between the Ed sensor at the top and the Lu sensor at the bottom is approximately 1.412 m. This presents a problem when defining a pressure co-ordinate system that is common to all sensors that make up the instrument. When processing optical data, the data closest to the surface is usually of greater importance than data at the deepest of the profile. Therefore it makes sense to choose the sensor nearest the surface which in all profiler cases is the Ed sensor. The pressure values for the Ed sensor, hereafter referred to as master pressure, are then used to derive the pressure co-ordinates. It should be noted that if no Ed sensor is present then the Lu sensor pressure must be used instead.

The master pressure is then filtered to eliminate frames where pressure is not increasing.

Equation 9 Pressure Filtering

depth(n+1)−depth(n)≤0

The pressure co-ordinates are created by rounding up the first master pressure value to the nearest 0.1 m then increasing the pressure by intervals equal to the depth resolution until reaching the highest (deepest) master pressure value which is rounded down to the nearest 0.1 m. For example:

| Master Pressure From Ed | Pressure Co-ordinates |
| --- | --- |
| 2.2657 | 2.3 |
| 2.3727 | 2.4 |
| 3.0684 | 2.5 |
| 4.2028 | 2.6 |
| 4.7808 | etc. |
| 5.4015 | 6.3 |
| 5.8510 | 6.4 |
| 6.3326 | 6.5 |
| 6.6323 | 6.6 |

#### Calculate Time Co-ordinates

Almost all reference instruments have an Es sensor which ProSoft uses to calculate the time co-ordinate system. In the case of SAS instruments, an Lt sensor is always present which is used to calculate the time co-ordinate system. There are three possible ways to calculate the time values, hereafter referred to as master time, for each frame of the Es or Lt sensor. They are listed below in order of precedence:

1. If a timer sensor is part of the instrument, then this is the most accurate way to calculate the master time. The master time for the very first frame is taken as the first Timetag2 value if available else Time Stamp, the time at which logging data begins, is used. The timer values give the elapsed time for each frame since the time of the first frame which is used to calculate the absolute time for each frame. Note that timers are available in Satnet systems as well as Hyperspectral systems.

2. If Timetag2 values were appended to the logged data then those values are used as the absolute time for each frame.

3. The absolute time for the first frame is taken as the Time Stamp. The rate for the sensor is then used to calculate the elapsed time for each frame which in turn is used to calculate the absolute time for each frame.

The time co-ordinates are created by taking the first master time value and then increasing it by an interval determined from the Es time data until reaching the last master time value. For Sas data the time interval is determined from the optical sensor operating at the fastest rate.

For example: (interval  = 0.100 sec)

| Master Time from Es | Time Co-ordinates |
| --- | --- |
| 17:57:41.167 | 17:57:41.167 |
| 17:57:41.333 | 17:57:41.267 |
| 17:57:41.500 | 17:57:41.367 |
| 17:57:41.667 | 17:57:41.467 |
| 17:57:41.833 | etc. |
| 17:57:42.000 | 17:57:42.167 |
| 17:57:42.167 | 17:57:42.267 |
| 17:57:42.333 | 17:57:42.367 |
| 17:57:42.500 | 17:57:42.467 |

### Coordinate Interpolation

The data to be averaged is interpolated onto the master co-ordinates ensuring that all the data is mapped to a common co-ordinate system prior to averaging. The co-ordinates are based on either pressure or time. The interpolation method used for both Profiler and Reference instruments is ‘linear’ and for SAS instruments ‘spline’ is used.

### Write Level 2s HDF File

The interpolated data is written to the corresponding level 2s HDF file.

### Wavelength Interpolation

Profiler, Reference and SAS data can be interpolated onto a constant wavelength interval by setting the level 3 parameter Wavelength Interpolation to the desired interval. Wavelength interpolation can only be used for Hyperspectral instruments. Multispectral instruments have too few wavelength channels to provide an accurate interpolation. It is highly recommended to use wavelength interpolation for Hyperspectral profilers/references as it will help maximize the number of matching wavelengths for Level 4 processing.

### Natural Log Transform

Optical sensor data (Lu, Es etc.) are natural log transformed to “straighten” the data prior to averaging.

### Average Data

The data is then averaged based on the level 3 parameters Bin Interval and Bin Width for pressure co-ordinates and Time Interval and Time Width for time co-ordinates. The Bin Interval or Time Interval determine the points at which to average with the Bin Width or Time Width determining the number of data points to use for averaging.

For example: BIN_INTERVAL = 0.5 m BIN_WIDTH = 0.3 m

Depth (m)

Ln(Lu)

0.9

1.0

1.1

1.2

1.3

1.4

1.5

1.6

1.7

1.8

1.9

2.0

2.1

2.2

2.3

-1.5265

-1.5530

-1.5722

-1.5932

-1.6164

-1.6411

-1.6650

-1.6870

-1.7124

-1.7386

-1.7581

-1.7806

-1.8087

-1.8287

-1.8487

BIN_INTERVAL = 1.0m

average = insufficient number of data points to average.

BIN_INTERVAL = 1.5m

average = -1.6648

BIN_INTERVAL = 2.0m

average = -1.7822

BIN_WIDTH 03.m

BIN_WIDTH 03.m

## Level 4 Data Processing

When using the level 4 processing menu a dialogue box allows the user to specify which level 4 products to include in the level 4 HDF file. Each option selected will be included as a separate sensor within the HDF file. The first two options, ‘Diffuse Attenuation Coefficient’ and ‘Propagate Optical Variables to Surface’ are sometimes needed to process other selections such as Water Leaving Radiance and will therefore be automatically selected when needed.

Notes On Wavelength Differences: When processing calculations that involve data from two different sources (e.g. comparing Lu and Ed sensors) the wavelengths from the two sources will never be exact (e.g. Lu 489.8nm and Ed 490.1nm), unless wavelength interpolation was selected at level 3a. Therefore ProSoft uses a matching tolerance between wavelengths in order for the channels to be used in the calculation. For example if Lu is 489.8nm and the closest Ed channel is 490.1nm (490.1 – 489.8 < tolerance) these two channels can be used together in a calculation. In Level 4 data presentation the wavelength shown is the mean value between the two wavelengths (e.g. 489.95nm from the above example). Tolerances are set as follows, multispectral 2 nm and hyperspectral 0.5 nm.

### Diffuse Attenuation Coefficient

Overview: Under typical conditions for which incident lighting is provided by sun and sky, the various radiance's and irradiances all decrease exponentially with depth, at least far enough below the surface (and far enough above the bottom in shallow water) to be free of boundary effects. This effect can be generally expressed as:

Equation 10 Diffuse Attenuation Coefficient (Ocean Optics Protocols for SeaWiFS eqn. 26 pg. 49)

Edz,λ=Ed0−,λe−0zkdz',λdz'

Edz,λ - downwelling irradiance as a function of depth and wavelength.

Ed0−,λ - surface value of downwelling irradiance as a function of wavelength.

e−0zkdz',λdz' - effect of the diffuse attenuation coefficient as a function of depth and wavelength.

The traditional method of k analysis (Smith and Baker 1984 and 1986) is to estimate k(z) as the local slope of measured ln(Ed(z)) in an interval of a few meters centered on depth zm.

Equation 11 Ocean Optics Protocols for SeaWiFS eqn. 31 pg. 50

lnEdz≃lnEd0−−z−zmkzm

The slope k and the intercept Ed(0-) are calculated using a least squares regression fit of the data over the number of regression points selected.

p = polyfit(x,y,1)  MatLab function.

The degree of the resulting polynomial is one.

For the shallowest and deepest depths of a profile there is not enough data points to meet the required number of points for regression. In this case the k value is calculated by selecting the required number of points nearest to the depth upon which the k value is centered. For example if 1m is the shallowest depth then the k value for 1m is based on level 3a data taken from 1m to 3m, if Integration Points = 5, and the k value for 2m is based on level 3a data taken from 1m to 4m. Therefore k values at the shallow and deep sections of the profile are unreliable and are not used for Level 4 data calculations.

#### Integration Points

1 m

3 m

5 m

2 m

4 m

6 m

7 m

8 m

Ed (uW/cm2/nm)

Depth

INTEGRATION_POINTS = 5

Slope = K(3m)

Ed(0-)

0 m

Slope = K(5m)

Note: Ed values are ln transformed prior to regression analysis which would straighten out the curve.

Note: Successive values of K are calculated using overlapping data points as can be seen in the above graph.

Figure 7 K Analysis

Integration points is the processing parameter that determines the number of data points to use for regression analysis. Only odd number values can be used (e.g. 3,5, 7 etc.) in order for depth z to be centered on the level 3a depth interval (e.g. 3m, 5m, 7m etc.).

In cases where there is only a reference instrument and no profiler, therefore no K_Lu, K_Ed or K_Eu, to propagate Ls to surface, then K is estimated from a time series of upwelling radiance measured at a fixed depth below the surface. In this case ProSoft estimates the spectrum of K(λ) in three steps:

- Following the Austin Petzold (1981) model, K is estimated at the wavelengths 443nm and 550nm.

Equation 12 Austin Petzold 490nm

K490=0.0833Lu443Lu550−1.491+0.022

Equation 13 Austin Petzold 520nm

K520=0.0663Lu443Lu550−1.398+0.044

These K values are used in step 2 to obtain the concentrations of pigments C at 490nm and 520nm.

- Morel’s (1988) empirical model for estimation of K is inverted to obtain the concentration of pigments (Chl a and Pheo):

Equation 14 Morel 1988

C=Kλ−Kwλχc1eλ

Kwλ - diffuse attenuation coefficient of water. Obtained from file MORELKW.FIT

χcλ - coefficient determined empirically. Obtained from file MORELCHI.FIT

eλ - coefficient determined empirically. Obtained from file MORELEPS.FIT

C is obtained as the average value of the two C’s calculated at 490nm and 520nm and is used in step 3.

- The spectrum of K(λ) is obtained using Morel’s equation 9 (1988):

Equation 15 Morel 1988 Equation 9

Kapλ=Kwλ+χλCeλ

To calculate Kap within 700-800nm, ProSoft uses the following steps:

i. The values of Kw(λ) within 700-800nm are obtained from Smith and Baker (1981) MORELKW.FIT.

ii. It is assumed that χ(λ) and e(λ) decrease linearly from the values given at χ(700) and e(700) to 0 and 800nm respectively.

iii. Using the above assumption and Kw within 700-800nm, K(λ) can be estimated at λ = 700-800nm interval.

The K values calculated depend on the instrument configuration. Possible values are:

K_Ed

K_Lu

K_Eu

Kap_Ls – K is calculated using the Austin Petzold model

### Propagate Optical Variables to Surface

The surface values of the profiler optical variables are obtained as the intercept of a least squares regression fit based on the shallowest depth centered on Zm.

Equation 16 Surface Variables (Ocean Optics Protocols for SeaWiFS eqn. 31 pg. 50)

lnEdz≃lnEd0−−z−zmKzm

The calculated intercept is a product of the diffuse attenuation coefficient (Kd) calculation. As a result, if optical variables are propagated to the surface, the Kd profile for each optical variable (Lu, Ed and Eu) will be calculated and included with the level 4 hdf file. In cases of shallow depths that do not meet the required number of  points for regression, the surface variable is calculated based on the intercept of the k slope from the first depth value that has the required number of points. For example in k analysis the surface optical variable Ed(0-) will be calculated on the intercept of the slope based on the k value at 3m and not at 1m or 2m.

Possible values include:

Profiler:

Ed(0-) calculated from K_Ed

Lu(0-) calculated from K_Lu

Eu(0-) calculated from K_Eu

Reference:

Es(0-) if Es is measured above the surface then Es(0-) is calculated using:

Equation 17 Es(0-)

Es0−=Esλ1+α

α – Fresnel reflection albedo for irradiance from sun and sky. Default value of 0.043 which can be changed as a level 4 parameter.

Ls(0-) is calculated using one of the following (whichever one is available starting with the top):

K_Lu

K_Ed

K_Eu

Kap_Ls – K is calculated using the Austin Petzold model

Example:

Equation 18 Ls(0-)

Ls(0−,λ)=Ls⋅ez⋅k

#### Reflection Albedo

Processing parameter that sets the value for the Fresnel reflection albedo for irradiance from sun and sky. A default value of 0.043 is used and can be changed as a level 4 processing parameter.

### Water Leaving Radiance

Normalized water leaving radiance is estimated with the following equation:

Equation 19 Normalized Water Leaving Radiance (Ocean Optics Protocols eqn. 63 pg. 54)

Lwn=LwλFoλEsλ

Fo(λ) – the mean extraterrestrial solar irradiance (Neckel and Labs 1984)

Es(λ) – downwelling spectral irradiance at z(0+)

Lw(0+,λ) – upwelling radiance propagated through the surface z(0+) using:

Equation 20 Water Leaving Radiance

Lw0+,λ=Lu0−,λ1−ρλ,θηw2λ

ρ(λ,θ) – Fresnel reflectance index of seawater. Default value of 0.021 for seawater which can be changed as a level 4 global variable.

ηw(λ) – Fresnel refractive index of seawater. Default value of 1.345 for seawater which can be changed as a level 4 global variable.

Es(λ) from water leaving radiance is usually obtained from an above water reference, but in cases where no reference data is available Ed(0-,λ), measured by a profiler, is used. Ed is extrapolated through the surface using:

Equation 21 Ed(0+)

Ed0+,λ=Ed0−,λ1+α

α – Fresnel reflection albedo for irradiance from sun and sky. Default value of 0.043 is used which can be changed as a level 4 processing parameter.

If Lu(0-,λ) from water leaving radiance is not available then Ls(0-,λ) is used instead. Ls(0-,λ) is calculated when propagating variables to the surface.

The values used for Fo(λ) (Neckel and Labs) can be found in the file Solar10.dat.

#### Reflectance Index

Fresnel reflective index for seawater. A default value of 0.021 is used and can be changed as a level 4 processing parameter.

#### Refractive Index

Fresnel refractive index for seawater. A default value of 1.345 is used and can be changed as a level 4 processing parameter.

### Surface Remote Sensing Reflectance's

Remote sensing reflectance's RL(0+,λ) are calculated using:

Equation 22 Surface Remote Sensing Reflectance

RL0+,λ=LW0+,λEd0+,λ

Ed(0+,λ) – denotes downwelling spectral irradiance measured just above the surface (Es) or extrapolated through the surface Ed(0+). For Es the median of the time series data is taken for each wavelength.

Lw(0+,λ) – denotes upwelling spectral radiance propagated through the surface.

The remote sensing reflectance's calculations produce several possible combinations depending on the instrument configuration. Possible values include:

i. for Lu and Ed sensors Lu(0+,λ)/Ed(0+,λ)

ii. for Lu and Es sensors Lu(0+,λ)/Es(λ)

iii. for Ls and Es sensors Ls(0+,λ)/Es(λ)

Ed is extrapolated through the surface using:

Equation 23 Above Surface Ed(0+)

Ed0+,λ=Ed0−,λ1+α

α – Fresnel reflection albedo for irradiance from sun and sky. A default value of 0.043 is used and can be changed as a level 4 processing parameter.

Lw is calculated using:

Equation 24 Water Leaving Radiance

Lw0+,λ=Lu0−,λ1−ρλ,θηw2λ

ρ(λ,θ) - Fresnel reflectance index of seawater. Default value of 0.021 is used.

ηw(λ) - Fresnel refractive index of seawater. Default value of 1.345 is used.

### Remote Sensing Reflectance Profile

Remote sensing reflectance profile is calculated as the ratio of upwelling radiance to downwelling irradiance with changing depth. Only Lu is used for upwelling radiance and Ed for downwelling irradiance.

Equation 25 Remote Sensing Reflectance Profile

RSR=Luz,λEdz,λ

### Surface Reflectance's

Surface reflectance is calculated as the ratio of upwelling irradiance just above the surface to downwelling irradiance just above the surface.

Upwelling irradiance can be derived from several sources depending on the instrument configuration and include Eu(0+), Lu(0+) converted to Eu(0+), Ev(0+) and Ls(0+) converted to Ev(0+).

Downwelling irradiance can be derived from several sources depending on instrument configuration and include Ed(0+) and Es.

Equation 26 Surface Reflectance

RFL=UpwellingIrradiance0+,λDownwellingIrradiance0+,λ

The following sensor combinations are possible:

i. Eu and Ed sensors: Eu and Ed are propagated above the surface using

Equation 27 Above Surface Ed(0+)

Ed0+,λ=Ed0−,λ1+α

α – surface albedo (default = 0.043)

RFL = Eu(0+,λ)/Ed(0+,λ)

ii. Lu and Ed sensors: Lu is propagated above the surface and converted to Eu using:

Equation 28 Above Surface Eu(0+)

Eu0+,λ=4.5⋅Lu0−,λ1−ρλ,θηw2λ

ρ – reflective index (default = 0.021)

ηw – refractive index (default = 1.345)

RFL = Eu(0+,λ)/Ed(0+,λ)

iii. Lu and Es sensors: Lu is propagated above the surface and converted to Eu using equation Eu(0+,).

RFL = Eu(0+,λ)/Es(λ)

iv. Ev and Es sensors : Ev is propagated above the surface using equation Ed(0+,).

RFL = Ev(0+,λ)/Es(λ)

v. Ls and Es sensors: Ls is propagated above the surface and converted to Ev using equation Eu(0+,).

RFL = Ev(0+,λ)/Es(λ)

### Reflectance Profile

The reflectance profile is calculated as the ratio of upwelling irradiance to downwelling irradiance. The most preferred method of calculation is to use upwelling irradiance Eu from the profiler and downwelling irradiance Ed, also from the profiler.

Equation 29 Reflectance Profile

RFL=Euz,λEdz,λ

### Photosynthetically Available Radiation

Photosynthetically available radiation is a measure of the number of photons available for photosynthesis by chlorophyll. Units of PAR are expressed as quanta/cm2s and is estimated using the following:

Equation 30 Photosynthetically Available Radiation

PAR=400nm700nmλhcEdλdλ

λ – wavelength

h – Planck’s constant 6.625x10-34 Js

c – speed of light 3x108 m/s

Ed(λ) – downwelling spectral irradiance (uW/cm2 /nm)

Prior to integration Ed(λ) is interpolated onto a constant 1nm interval and subject to the following rules:

if λ1 > 400nm then Ed(400) = Ed(λ1)

if λn < 700nm then Ed(700) = Ed(λn)

Included with the PAR data is the percentage of PAR for various depths. This ratio is calculated as:

Equation 31 Percentage PAR

%PAR=PARzPAR0+

The attenuation coefficient of PAR (K_PAR(0-)) is also calculated near the surface using the same methods as described under the diffuse attenuation coefficient. This value is included in the data table attributes as K_PAR(0-).

If reference Es data is available then PAR is calculated as:

Equation 32 Reference Par

PAR=400nm700nmλhcEsλdλ

The median value of PAR is then calculated and included in the vdata table attributes and called Reference PAR.

### Chlorophyll a Profile Estimates Morel 98 Model

The estimation of chlorophyll a concentration (ug/l) in profile is estimated for each depth level given from the level 3a hdf file. The method used is as follows:

Equation 33 Morel 98 Chlorophyll Model

C=Kλ−KWλχCλ1eλ

χC, KW and e(λ) are taken from the Morel 98 model. These values can be seen in the files MOR98CHI.DAT, MOR98KW.DAT and MOR98EPS.DAT located in the ProSoft installation directory.

The Morel 98 model is derived from equation 9 pg. 10753 from:

“Optical Modeling of the Upper Ocean in Relation to Its Biogenous Matter Content”

Andre Morel

Journal of Geophysical Research Volume 93 No. C9, pgs. 10749-10768

### Chlorophyll a Surface Estimates SeaBAM OC2 Model

Chlorophyll a concentration is estimated at the water surface using the remote sensing reflectance values calculated previously.

Equation 34 Chlorophyll a Surface Estimate SeaBAM OC2 Model

C=100.2974−2.2429R+0.8358R2−0.0077R3−0.0929

Equation 35 Calculation of R Coefficient

R=Log10Rrs490nmRrs555nm

The calculation of R coefficient uses the remote sensing reflectance values at 490nm and 555nm to calculate the value of R. Therefore several different calculations of Chlorophyll a are possible depending on how the remote sensing reflectance was calculated. The possible combinations are Lu(0+)/Es, Lu(0+)/Ed(0+) and Ls(0+)/Es.

The value of R is then used in to calculate the chlorophyll a concentration (ug/l) at the surface.

Equations for chlorophyll a surface estimate and calculation of R coefficient are derived from:

“Ocean Color Chlorophyll Algorithms for SeaWiFS”

O’Reilly et al.

Journal of Geophysical Research, Volume 103, No. C11, pgs. 24937-24953

### Chlorophyll a Surface Estimates Gordon 88 Model

Chlorophyll a concentration is estimated at the water surface using the normalized water leaving radiance's calculated previously.

Equation 36 Chlorophyll a Surface Estimate Gordon 88 Model

C=1.15Lwn443nmLwn560nm−1.42C≤1μgl

Equation 37 Chlorophyll a Surface Estimate Gordon 88 Model

C=3.64Lwn500nmLwn560nm−2.62C>1μgl

The equations for Chlorophyll a surface estimate use normalized water leaving radiance's values at 443nm, 500nm and 560nm. Therefore several different calculations of chlorophyll a are possible depending on how the normalized water leaving radiance was calculated. The possible combinations are Lu(0+)/Ed(0+), Lu(0+)/Es and Ls(0+)/Es.

Equations for Chlorophyll a surface estimates are taken from:

“Ocean Color Remote Sensing Systems: Radiometric Requirements”

H. Gordon

Processing for Remote Sensing, pgs. 151-167

### Estimation of Energy Fluxes

Energy fluxes are estimated using the following equation:

Equation 38 Estimation of Energy Fluxes

F=100400700Edλdλ

ProSoft uses Matlab’s trapezoidal numerical integration routine ‘trapz.m’ to perform the integration. Prior to integration, a minimum-maximum wavelength check is performed then the irradiance data is interpolated onto a constant 1nm interval. Since irradiance data is measured in units of μWcm-2nm-1, a scaling factor of 100 is used to convert from cm-2 to m-2.

# Appendix A: Term Definitions

## File Naming Convention

HDF File – ProSoft input or output data file where meta-information and data have been organized using Hierarchical Data Format following the principles of data organization and processing.

Instrument File – a calibration (CAL) file or a telemetry definition file (TDF) ( for details see Satlantic Instrument File Standard).

File Nametag – last letters or numbers which follow underscore ('_') in a file name.

Level 1 – binary logged data file with the RAW extension.

Level 1a – binary data is extracted from RAW data under the control of the instrument (calibration) files. Extracted information is grouped along with its calibration information and placed into Level la hdf files. File nametag is _L1a.

Level 1b – Level 1a data is calibrated. No data editing is applied. Shutter darks, if present, are not applied. File nametag is _L1b.

Level 2 – includes Level 1b data, which is further modified per request basis. File nametag is _L2.

- Shutter dark correction is applied; reference and dark data deglitching is applied.

- Profiler's data is tilt edited

- For SAS data, additionally to original collected data, a new data group is generated, which includes the interpolated lookup tables of GPS co-ordinates, time, optical, and meteorological data.

Level 2s – level 2 data is interpolated onto a common co-ordinate vector, which is either depth (Profiler) or time (Reference only or Sas). File nametag is _L2s.

Level 3a – Averages level 2s data as defined by level 3 processing parameters

Level 4 – Data products produced from level 3a data.

## Optical Sensors

Ed - Profiler downwelling plane irradiance

Lu - Profiler upwelling radiance

Eu - Profiler upwelling plane irradiance

Ld - Profiler downwelling radiance (special applications only)

Es - Reference downwelling plane irradiance

Ls - Reference upwelling radiance

Ev - Reference upwelling plane irradiance

Ef - Reference downwelling plane diffuse irradiance (special applications only)

Li - SAS sky radiance

Lt – SAS water leaving radiance

Units: Irradiance - uW/cm2/nm

Radiance - uW/cm2/nm/sr

## Ancillary Sensors

Tw     - External water temperature (oC)

Ti     - Internal irradiance sensor array temperature (oC)

Tr     - Internal radiance sensor array temperature (oC)

Pres   - Profiler depth (m)

TiltX  - Sensor tilt on X axis (deg)

TiltY  - Sensor tilt on Y axis (deg)

Rtilt  - Reference tilt about vertical axis (deg)

Ptilt  - Profiler tilt about vertical axis (deg)

Pitch  - Sensor tilt along horizontal axis (deg)

Roll   - Sensor tilt about horizontal axis (deg)

Cond   - Seawater conductivity (mmho/cm)

Fluor  - Fluorometer (ug/l)

Trans  - Transmissometer (/m)

Comp   - Flux gate compass (deg)

Altim  - Altimeter (distance from bottom) (m)

PAR    - PAR sensor (uE/m2)

Time   - system time (decimal hours)

Vel    - profiler descent velocity (m/sec)

Frame  - Instrument frame counter (0..255)

LuDark - Profiler radiance dark detector OCR-1000 only (counts)

EdDark - Profiler irradiance dark detector OCI-1000 only (counts)

EsDark - Reference irradiance dark detector OCI-1000 only (counts)

LsDark - Reference radiance dark detector OCR-1000 only (counts)

# Appendix B: Data Formats

Table 8 File Header

| File Header | Units | Description |
| --- | --- | --- |
| PROSOFT | None | ProSoft software version used in analysis |
| CRUISE-ID | None | Identifies the cruise on which the current series of log files is taken. |
| OPERATOR | None | Identifies the operator of the data acquisition software for the current log file. |
| LATITUDE | Decimal Degrees | Identifies the latitude coordinates of the instrument for the current log file. |
| LONGITUDE | Decimal Degrees | Identifies the longitude coordinates of the instrument for the current log file. |
| ZONE | None | Identifies the data acquisition time zone of the current log. |
| COMMENT | None | Comment field for ancillary sky and sea observations |
| STATION-ID | None | Identifies the station for the current series of log files. |
| CAST | None | Identifies the current log file in a series of related log files. |
| TIME-STAMP | None | Identifies the time on the data acquisition computer as of the beginning of the current log file. |
| MODE | None | Identifies the data acquisition software mode. |
| PRESSURE-TARE | Meters (m) | Pressure tare taken at the surface for current log file. |
| TIMETAG | None | Obsolete. |
| DATETAG | None | Switch to include date/time tagging. |
| TIMETAG2 | None | Switch to include timetag2 tagging. |
| PROFILER | None | Identifies the status of the profiler instrument. |
| REFERENCE | None | Identifies the status of the reference instrument. |
| PRO-DARK | None | Profiler dark correction method. |
| REF-DARK | None | Reference dark correction method. |
| PRO-ID | None | This field is the identifier for the profiler. |
| PROCESSING_LEVEL | None | Identifies the processing level of the current file. |
| FILE_CREATION_TIME | None | Date of file creation. |
| DEGLITCH_PRODAT | None | Switch to include deglitching of profiler data. |
| DEGLITCH_REFDAT | None | Switch to include deglitching of reference data. |
| PRESS_TARE_CORRECTION | Meters (m) | Pressure tare value applied to current log file. |
| DEPTH_RESOLUTION | Meters (m) | Depth resolution used for Level 2s data interpolation. |
| BIN_INTERVAL | Meters (m) | Level 3a variable that determines which depth interval to average. |
| BIN_WIDTH | Meters (m) | Level 3a variable that determines depth range within which to average. |
| TIME_INTERVAL | Seconds | Level 3a variable that determines which time interval to average. |
| TIME_WIDTH | Seconds | Level 3a variable that determines the time range within which to average. |
| WAVEL_INTERP | (nm) | Level 4 variable that determines wavelength interpolation interval. |
| INTEGRATION_POINTS | None | Level 4 variable that determines number of data points to use for regression analysis. |
| REFLECTION_ALBEDO | None | Reflection albedo value used in level 4 processing. |
| REFLECTANCE_INDEX | None | Reflectance index value used in level 4 processing. |
| REFRACTIVE_INDEX | None | Refractive index value used in level 4 processing. |

Table 9 Level 1-2 Table Descriptors

| Table Label | Units | Description |
| --- | --- | --- |
| PRES | Meters (m) | Used to define a pressure sensor measurement. Pressure measurements are usually used with submersible instruments to gauge their depth. |
| ALTIM | Meters (m) | Used to define an altimeter sensor measurement. Altimeter measurement is normally used to gauge the distance from a submersible instrument to the ocean floor. |
| FRAMECOUNTER | None | The value of a frame counter is simply a numerical counter that can help ensure all frames received are sequential. |
| CHECKSUM | None | The value of the check sum is the last byte of the sum of all bytes, up to but not including the check sum sensor, in a frame of telemetry subtracted from 0. |
| INTTIME | Seconds | Defines the adaptive or adjusted integration time for all subsequent optical sensors for which the sensor was defined. |
| DARK | None | Some optical sensors are shielded from light exposure to establish a dark noise floor for a related optical sensor grouping. |
| DARK_SAMP | None | Indicates the number of dark sensors used to calculate a dark average. |
| DARK_AVE | None | Defines the actual average dark value. |
| T | Degrees Celsius (C) | Used to define a temperature sensor measurement. Temperature sensors can be placed anywhere on an instrument. |
| TILT | Degrees | Used to define a tilt or attitude measurement for an instrument. A tilt measurement is normally used to determine the alignment of an instrument with the vertical. |
| PITCH | Degrees | Used to define a tilt sensor along the horizontal axis. A pitch sensor can be used to determine an instruments orientation relative to the horizon. |
| ROLL | Degrees | Used to define a tilt sensor about the horizontal axis. A roll sensor can be used to determine an instruments orientation relative to the horizon. |
| COMP | Degrees | Used to define a flux gate compass sensor measurement. These sensors measure magnetic flux to establish a compass direction. |
| COND | (mmho/cm) | Used to define a seawater conductivity sensor measurement. A conductivity measurement is necessary for calculation of salinity. |
| FLUOR | (ug/l) | Used to define a Fluorometer sensor measurement. |
| TRANS | (/m) | Used to define a Transmissometer sensor measurement. A Transmissometer measurement indicates the attenuation of a light beam in a given media. |
| PAR | (E/m2) | Used to define a sensor for measuring photosynthetically active radiation. PAR measurements indicate the total light energy available for photosynthesis per unit area. |
| TIMER | Seconds | Used to define a timer sensor for indicating the number of seconds that have elapsed since the instrument was powered up. |

Table 10 GPS Table Descriptors

| Table Label | Units | Description |
| --- | --- | --- |
| UTCPOS | None | Used to define the Universal Coordinated Time of the GPS frames |
| LATPOS | Degrees | Used to define the latitude coordinates of the GPS frames. |
| LONPOS | Degrees | Used to define the longitude coordinates of the GPS frames. |
| LATHEMI | None | Used to define the hemisphere of the associated latitude coordinates. |
| LONHEMI | None | Used to define the hemisphere of the associated longitude coordinates. |
| SPEED | Knots | Used to define the ground speed of the GPS frames. |
| COURSE | Degrees | Used to define the course over the ground in degrees true. |
| MAGVAR | Degrees | Used to define the magnitude of the magnetic variation to be used between true and magnetic degrees. |
| DATE | None | Used to define the date of the GPS measurements (ddmmyy) |
| MODE | None | Used to define the position system mode for the GPS sensor (i.e. A = Autonomous, D = Differential) |

Table 11 Level 4 Table Descriptors

| Table Label | Units | Description |
| --- | --- | --- |
| K_Lu , K_Ed, K_Eu | (/m) | K values |
| Lu(0-), Ed(0-), Eu(0-), Es(0-) | (w/cm2/nm) or (w/cm2/nm/sr) | Surface optical data |
| Lw-Lu(0+), Lw-Ls(0+) | (w/cm2/nm/sr) | Water leaving radiance. |
| Lwn-Lu(0+)/Ed(0+), Lwn-Lu(0+)/Es, Lwn-Ls(0+)/Es | (w/cm2/nm) | Normalized water leaving radiance. |
| RSR-Lu(0+)/Ed(0+), RSR-Lu(0+)/Es, RSR-Ls(0+)/Es | (/sr) | Surface remote sensing reflectance. |
| RSR-(Lu/Ed) | (/sr) | Remote sensing reflectance profiles. |
| RFL-Eu(0+)/Ed(0+), RFL-Eu(0+)/Es | (/sr) | Surface reflectances. |
| Profiler Par | (quanta/cm2s) | Photosynthetically available radiation |
| Morel 98 Chlorophyll Profile, SeaBAM OC2 Surface Chlorophyll, Gordon 88 Surface Chlorophyll | (g/l) | Chlorophyll Estimates |
| Profiler Flux, Reference Flux | (w/m2) | Energy fluxes |
| Water Properties Table | Depth (m) Temp (C) Conductivity (mmho/cm) Conductivity Ratio (None) Salinity (psu) Density (Kg/m3) | Water Properties |

# Appendix C: Profiler/Reference Data Deglitching Function

## Background

Deglitching of optical data measured in sea environment is not a trivial task. Data glitches that can be caused by power instabilities, inconsistencies between data transfer baud rates and data logging speed, and other reasons that are not caused by changes in the natural environment. On the other hand, noise in optical data can be also be caused by wave-focusing effects, random changes in profiler's tilts, and the noise caused by dark current.

We make difference between the data glitches and remaining part of unwanted noise mainly to make possible differential data treatment depending on the cause of a noise. In a case of glitches, we want the respective part of data simply removed. In a case of wave focusing, dark current (DC) influence and profiler tilts, one may try to reduce their influence by some sort of data filtering. In an extreme case, data with significant tilts, DC etc can be also removed.

## Description of data distribution with depth

1. Profiler's optical data changes exponentially with depth. Changes in signal values can be as large as 8 orders e.g. from 1 to 1e-8.

2. Part of the data changes that can be regarded as noise, also change significantly. Moreover, noise distribution with depth is not homogenous. The causes for noise in the different parts of a profiler are different e.g. if in the surface layer wave focussing and tilt effects are dominating then with the depth increase, dark current noise contribution increases, and finally starts to dominate.

## Problem

To identify glitches, one should divide measured data profile (S) into two parts: signal (<S>) and noise (E).

S=S+E						(1)

where operator <> means some sort of data averaging. From equation (1) it is clear that noise definition E=S−S depends on a meaning of operation of data averaging. The measurements with E > Threshold will be declared glitches. Mean signal <S> can be estimated using various approaches. E.g. FFT, wavelet transform, measured data fit with a semi-empirical or empirical model (e.g. simple exponential or polynomial fit), simple moving average etc. In all these cases, distribution of E with depth will remain statistically inhomogeneous i.e. one could expect that p.d.f of E will change as a function of depth. The outlined inhomogeneous error distribution constitutes the main problem for glitches identification and their subsequent removal.

## Approach for the Problem Solution

1. We assume that we have data S measured at fixed time or depth interval. For simplicity we will use data frame number 'i' instead of a time or space co-ordinate. Thus we have

S(i), where i = 1,…, N, where N is number of measurements

2. From S(i) we calculate the ratio

dS(i) = S(i+1)/S(i), i = 1, …, N-1

3. Sort dS in an ascending order

dS_sorted = sort(dS)

4. Calculate a measure of data deviation

stdS = dS_sorted(n1) - dS_sorted(n2), where

n1 = 0.2*nn

n2 = 0.75*nn

nn = length of dS

5. Calculate median value of dS

medN = median(dS)

6. Find data invalid data values

|dS(i) – medN| > NoiseThresh*stdS

The default values for a profiler and a reference instrument are 5 and 20 respectively. They should be adjusted according to a concrete situation.

From Step 6 it follows that all values that lie above NoiseThresh*stdS, should be removed.

## Examples

Figures 1 and 2 show the examples of measured upwelling radiance signal Lu without glitches (Figure 5b) and with glitches (Figure 6b). The position of the vertical red lines in the figures 1a, 1c and 2a indicate the levels below or above which data will be declared as a glitch.  Figure 5 shows data that had not any glitches in a sense defined above. However, one can see that significant part of data dominated by dark noise at the bottom of the profile is declared as the glitches. Although, usually this is not a problem to reject data dominated by DC, the given example demonstrates a potential problem in making difference between different causes of noise. In some cases wave focussing effects (practically not visible in the present examples) are also clearly visible in the upper part of a profile, and can be easily misinterpreted by a program as glitches in data. For this reason, sometimes it could be advantageous to perform data deglitching analysis only to some parts of a profile by identifying upper and lower depth boundaries of a profile. The deglitching function appended to this document enables a user to define the respective boundaries using ProSoft’s configuration file (see the last section of this note). In this way, depending on an aim of deglitching analysis, one could also remove or not touch significant amount of data affected dark noise or wave focusing.

Figure 8 Distribution of upwelling Lu radiance

with depth with no data glitches. (a) Distribution ratio of Lu(i+1)/Lu(i), two vertical red lines stay for medN| +/- NoiseThresh*stdS, (b) distribution of Lu as a function of pressure, (c) and (d) – the same as (a) and (d) but for downwelling irradiance.

Figure 9 The same as Figure 5 (a) and (b) but with apparent data glitches. (c) distribution of deglitched Lu as a function of depth.