Calibration of WETLabs’ Eco-type VSF meters

Emmanuel Boss

School of Marine Sciences

University of Maine

12/19/2005

Materials: WETLabs Eco-series backscattering sensor.

Mie code.

Weighting function of instrument response with respect to angle (W1()) and wavelength (W2()).

Calibration beads of known size distribution and optical properties.

Methods:

Measure the dark current of the instrument by placing the instrument in water with a black tape masking the detector.

Measure the signal in DIW in a black bucket.

Add beads in a logarithmic series (e.g. 2-4-8-16-32-64 etc’) and in each case record the counts from the VSF meter as well as beam attenuation and absorption with the ac-9/ac-s.

Using the Mie code generate the VSF,βλ,θ,D as function of wavelength (), diameter (D) and angle (), and the associated efficiency factor Qbλ,D for the wavelength(s) of the sensor.  The following integral provides the weighted (by angle, wavelength and size) angular distribution function:

βλ0,θ0=λminλmax02πrminrmaxW2θW1λNDβθ,λ,DdDsinθdθdλ1402πW2θsinθdθλminλmaxW1λrminrmaxNDπD2Qbλ,DdDdλ 	(1)

Where W1() is the wavelength weighting function (e.g. A Gaussian with a known FWHM), W2() the angular weighting function of the sensor, and N(r) the particulate size distribution of the beads. Given that these three function appear in both nominator and denominator, their absolute magnitude is not needed, only their shape. The shape of W1() is given based on the source output spectra, the spectrum of the filter on the receiver side and the photodector’s sensitivity. The shape of W2() from WETlabs. N(r) is given from the beads manufacturer. We obtain the needed phase function by normalizing  βλ0,θ0 (e.g. see below).

Rather then solve (1) directly we perform it in a series of steps:

1. Compute a size and wavelength weighted angular scattering distribution function:

βλ0,θ=k=1Kj=1JW1λjNDkβjkλj,Dk,θ.				(2)

We discretized the wavelength weighting function, W1(), to a 100 wavelenth values (J=100) spanning in equal intervals three standard deviations on each side of the wavelength of maximal value for the empirical curve describing W1()  (see appendix).  Similarly we discretized the calibration beads size distribution function to a 100 wavelenth values (K=100) spanning in equal intervals three standard deviations on each side of the modal size for the curve describing N(r) (see appendix).

2. Compute the associated phase function by requiring that its spatial integral be 1:

βλ0,θ=βλ0,θ2πφ0πβλ0,θsinθdθ					(3)

The limit of the integral in the denominator, 0, accounts for the fact that when we compute scattering we do it with an instrument of a finite acceptance angle (0.7 degree for the ac-9 or ac-S used here).

3. Convolve the phase function with the angular responsivity function to obtain the angle averaged phase function at 0:

βtheoryλ0,θ0=2πθ0πW2θβλ0,θsinθdθ2πθ0πW2θsinθdθ				(4)

For a given set of calibration beads and a given instrument (assuming no changes in spectral output or filters)βtheoryλ0,θ0 is constant. We proceed with the calibration procedure by performing a series of measurements of beads of known properties with both the VSF meter and an ac-9. For non-absorbing beads, cp can substitute for bp (e.g. a single band beam transmissometer can be used for the calibration).

The signal measured for k dilution experiments relates to the ac-9 measurements as follows:

signalk=darksignal+Aβtheoryλ0,θ0bp,kexp−ckr 		(2)

Where bp,k is the particulate scattering coefficient for dilution k, ck the associated total beam attenuation (this attenuation correction is most often negligible and account for attenuation along the path).

Regression analysis using the dilution series yields A, the factor needed to obtain the VSF in future measurements:

βλ0,θ0=signal−darksignalexpcrA

Why should wavelength be important?

The solution of scattering of light by a spherical particle was solved by Mie (1905). It can be showed that the solution depends on the index of refraction and on a size parameter, x=n, where D is the particle diameter, n the index of refraction of the medium and  the wavelength relative to the same medium the index of refraction is referenced to.

Similar changes in x will yield similar changes in the output of the solution thus:

dxx=dλλ2+dDD2+dnn2

Typically, for the sensors we analyzed here, dn/n<1% in the range of wavelength response. For the instruments analyzed here and the beads used by WETLabs (D=2m dtdev= 0.08m), dn/n~1%, dD/D=0.08/2=4%, and d/~20/500=4%. These result I changes in x (compared to accounting for only size changes, as currently performed by WETLabs) that can result in significant changes in βtheoryλ0,θ0, as shown below.

Results

I compared the values of βtheoryλ0,θ0 with those generated by WETLabs for beads with median size of 2m and standared deviation of 0.08m (Table 1) for the bb-9 specs I measured or obtained from WETLabs (see appendix, source output curves measured using Satlantic HyperPro).

Discrepencies with values up to 7% are found between the values currently used by WETLabs and those calculated here. When I compute βtheoryλ0,θ0for a hypothetical instrument with a 1nm FWHM around the nominal wavelength the discrepancies with WETLabs values drop to less than 1.6%. highlighting the importance of wavelength averaging.

|  | max | FWHM | Factors WETLABS | Factors MISC | Factors mono | relative difference MISC-WETLabs | Relative difference Mono-WETLabs |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 412 | 406.5 | 14.5 | 0.009937765 | 0.0099462 | 0.0100334 | 0.000848053 | -0.009577298 |
| 440 | 439.5 | 13 | 0.010592302 | 0.0108131 | 0.0106535 | 0.020419491 | -0.00576095 |
| 488 | 485.5 | 16.5 | 0.008712004 | 0.0090828 | 0.0086855 | 0.04082403 | 0.003046816 |
| 510 | 510 | 21 | 0.007381331 | 0.0074913 | 0.0073629 | 0.014679615 | 0.002500042 |
| 532 | 525.5 | 16 | 0.006256404 | 0.0067074 | 0.0062431 | 0.067238647 | 0.002128645 |
| 595 | 592.5 | 14 | 0.004166 | 0.0042296 | 0.0041128 | 0.015036883 | 0.012852104 |
| 650 | 651.5 | 14.5 | 0.003197367 | 0.0032285 | 0.003179 | 0.009643116 | 0.005761023 |
| 676 | 679 | 27 | 0.002929816 | 0.0029209 | 0.0029142 | -0.003052518 | 0.005344304 |
| 715 | 720.5 | 18 | 0.00260853 | 0.0025842 | 0.0025657 | -0.009414829 | 0.016555044 |

Table 1. Published nominal wavelength, wavelength of maximal response, FWHM of wavelength responsivity function W1(),βtheoryλ0,θ0 currently used by WETLabs, βtheoryλ0,θ0 determined using the approach presented here, βtheoryλ0,θ0 determined using the approach presented here assuming the , W1()as FWHM=1nm around the nominal wavelengths, the relative differences between the approaches.

For the Eco-VSF sensor the results are also different when wavelength characteristics are taken into account. In particular, for the 532 and 660nm nominal wavelengths, the differences (a bias!) seem to increase with the angle to a maximum of 5%.

|  | Centroid angle | Factors WETLABS | Factors MISC | relative difference MISC-WETLabs |
| --- | --- | --- | --- | --- |
| 660 | 100 | 0.00316691 | 0.0032835 | 0.035506533 |
| 660 | 125 | 0.0029804 | 0.0031074 | 0.040870052 |
| 660 | 150 | 0.00202236 | 0.00211 | 0.041537299 |
| 528 | 100 | 0.005728 | 0.0059232 | 0.032912902 |
| 528 | 125 | 0.006448 | 0.0066665 | 0.032834936 |
| 528 | 150 | 0.004504 | 0.0047325 | 0.048207142 |
| 440 | 100 | 0.007680 | 0.007754 | 0.009531364 |
| 440 | 125 | 0.011368 | 0.0114483 | 0.006985317 |
| 440 | 150 | 0.009426 | 0.0094099 | -0.001728467 |

WETLabs, βtheoryλ0,θ0 determined using the approach presented here, and the relative difference between the approaches.

Summary

It seems obvious from the results presented here that WETLabs needs to take into account the wavelengths characteristics of its sensors. At least a few in every series should have their source output and receiver responsivity be fully characterized (including effects due to epoxy and windows (if present). Similarly, the angular response of the sensors needs to be evaluated as well (e.g. Fig. A3) to establish the uncertainties of the calibration coefficient contributed by the uncertainties inβtheoryλ0,θ0.

Appendix

The different weighting functions needed to evaluate equ. 1. We are interested in their ‘shape’ not in their magnitude (as we normalize by it).

- W1()- this weighting function depends on the light source and receiver (filters are inserted before the receiver electronics). Three weighting functions contribute to W1(): The output of the LED (we measured it with a calibrated radiometer and found some notable differences with those provided by WETLabs, Fig. A1), the spectral characteristics of the  filter in front of the detector (obtained from the manufacturer) and the detector (photodetector) responsivity itself (obtained from the http://beammeasurement.mellesgriot.com/tut_photo_det.asp). W1() is the product of these three functions (Fig. A2).

Figure A1. Normalized source output for MISC’s bb9 (solid line) vs. that provided by WETLabs (dashed line).

Figure A2. Normalized source output (solid line, measured at MISC), detector filter sensitivity (dotted line, provided by WETLabs), and photodiode responsiivity (dashed line, found on WWW). The shape of W1() for each sensor is the product of these three curves.

- W2()- this weighting function depends on the geometry of the light source and receiver. It is the product of the angular response function, of the source and receiver, i.e. the likelihood that a photon will be emitted by the detector will make it into the receiver with scattering angle  assuming it had backscattered. We obtain it from the manufacturer (Fig. A3).

Figure A3. Normalized angular response. The shape of W2() was provided by WETLabs.