## 440(blue)
PO Box 518 | [5 empty] | (541) 929-5650
620 Applegate St. | [5 empty] | Fax (541) 929-5277
Philomath, OR 97370 | [6 empty] | www.wetlabs.com
Scattering Meter Calibration Sheet
2004-06-29 07:47 |  | Customer: | Emmanual Boss
Wavelength: 440 |  | S/N#: | BB9-132 |  | Job #: | 404020 | Tech: | K.C.
Use the following equation to obtain "scaled" output values:
b(qc) m-1 sr-1 = Scale Factor x (Output - Dark Counts)
• | Scale Factor for 440 nm |  | = | 4.7117e-05 | (counts)
• | Output |  | = | meter output | (counts)
• | Dark Counts |  | = | 52.818 | (counts)
| Instrument Resolution |  | = | 2.158 | (counts) | 0.00010168 | (m-1 sr-1)
Definitions:
• Scale Factor: Calibration scale factor, b(qc)/counts. Refer to User's Guide for derivation.
• Output: Measured signal output of the scattering meter.
• Dark Counts: Signal obtained by covering detector with black tape and submersing sensor in water.
Instrument Resolution: Standard deviation of 1 minute of collected data.
1. For each scattering centroid angle (qc), we compute the weighting function W(q,qc), by numerical integration of sample volume elements according to the sensor geometry.
2. We determine scattering phase functions, b(q,l)/b(l), for the microsphere calibration particles by weighing volume scattering functions computed from Mie theory according to the known size distribution of the microsphere polydispersion and normalizing to total scattering.
3. By convolving W(q,qc) with b(q,l)/b(l), we then compute the normalized volume scattering coefficient for each measurement angle, b(q,l)/b(l), with units of 1/sr: b(qc)/b for 2.00 micron diameter beads.
Wavelength, l (nm)
|  | Angle, qc | 440
|  | 117 | 0.010594
4. We then experimentally obtain raw scattering counts simultaneously with attenuation coefficients (Cp, using an ac-9) for a concentration series of the bead polydispersion. Absorption by the calibration particles is assumed negligible.
Results for this instrument follow:
raw counts
|  | Cp440 | 117
|  |  | 52.818 | Dark Counts: obtained by covering detector with black tape and submersing sensor in water.
| Clean water | 0.003 | 63.323
|  | 0.154 | 93.865
|  | 0.267 | 122.235
|  | 0.465 | 167.405
|  | 0.762 | 239.212
|  | 1.165 | 326.917
|  | 1.953 | 498.816
5. We obtain b/counts from the slope of a linear regression between Cp (equivalent to b for the beads) and counts:
117
| slope, intercept | 0.0044478 | -0.27929
|  | 3.8485e-05 | 0.0099584
| r | 0.99963 | 0.014513
6. Multiplying the experimental b/counts by the theoretical b(qc)/b yields the calibration scaling factor, S:
S
| 117 | 4.7117e-05
To obtain b(qc), 1) subtract the dark counts (see step 4) from the raw counts measured, then 2) multiply by S from the table above.
7. The test also provides a measure of the inherent noise level of the instrument from the short term RMS deviation from the average number of counts. This translates into the resolution of b(qc) (minimum detectable signal change) in units of m-1 sr-1).
| Noise (counts) |  | Resolution (1/[m. ster])
| 117 | 2.158 |  | 0.00010168

## 488(nm)
PO Box 518 | [5 empty] | (541) 929-5650
620 Applegate St. | [5 empty] | Fax (541) 929-5277
Philomath, OR 97370 | [6 empty] | www.wetlabs.com
Scattering Meter Calibration Sheet
2004-06-29 07:48 |  | Customer: | Emmanual Boss
Wavelength: 488 |  | S/N#: | BB9-132 |  | Job #: | 404020 | Tech: | K.C.
Use the following equation to obtain "scaled" output values:
b(qc) m-1 sr-1 = Scale Factor x (Output - Dark Counts)
• | Scale Factor for 488 nm |  | = | 1.1104e-05 | (counts)
• | Output |  | = | meter output | (counts)
• | Dark Counts |  | = | 41.606 | (counts)
| Instrument Resolution |  | = | 2.4543 | (counts) | 2.7252e-05 | (m-1 sr-1)
Definitions:
• Scale Factor: Calibration scale factor, b(qc)/counts. Refer to User's Guide for derivation.
• Output: Measured signal output of the scattering meter.
• Dark Counts: Signal obtained by covering detector with black tape and submersing sensor in water.
Instrument Resolution: Standard deviation of 1 minute of collected data.
1. For each scattering centroid angle (qc), we compute the weighting function W(q,qc), by numerical integration of sample volume elements according to the sensor geometry.
2. We determine scattering phase functions, b(q,l)/b(l), for the microsphere calibration particles by weighing volume scattering functions computed from Mie theory according to the known size distribution of the microsphere polydispersion and normalizing to total scattering.
3. By convolving W(q,qc) with b(q,l)/b(l), we then compute the normalized volume scattering coefficient for each measurement angle, b(q,l)/b(l), with units of 1/sr: b(qc)/b for 2.00 micron diameter beads.
Wavelength, l (nm)
|  | Angle, qc | 488
|  | 117 | 0.008712
4. We then experimentally obtain raw scattering counts simultaneously with attenuation coefficients (Cp, using an ac-9) for a concentration series of the bead polydispersion. Absorption by the calibration particles is assumed negligible.
Results for this instrument follow:
raw counts
|  | Cp488 | 117
|  |  | 41.606 | Dark Counts: obtained by covering detector with black tape and submersing sensor in water.
| Clean water | 0.004 | 63.968
|  | 0.174 | 193.757
|  | 0.304 | 296.206
|  | 0.535 | 477.946
|  | 0.901 | 768
|  | 1.344 | 1107.389
|  | 1.765 | 1464.429
|  | 2.25 | 1812.079
5. We obtain b/counts from the slope of a linear regression between Cp (equivalent to b for the beads) and counts:
117
| slope, intercept | 0.0012745 | -0.075542
|  | 7.7447e-06 | 0.0075497
| r | 0.99978 | 0.013011
6. Multiplying the experimental b/counts by the theoretical b(qc)/b yields the calibration scaling factor, S:
S
| 117 | 1.1104e-05
To obtain b(qc), 1) subtract the dark counts (see step 4) from the raw counts measured, then 2) multiply by S from the table above.
7. The test also provides a measure of the inherent noise level of the instrument from the short term RMS deviation from the average number of counts. This translates into the resolution of b(qc) (minimum detectable signal change) in units of m-1 sr-1).
| Noise (counts) |  | Resolution (1/[m. ster])
| 117 | 2.4543 |  | 2.7252e-05

## 555(nm)
PO Box 518 | [5 empty] | (541) 929-5650
620 Applegate St. | [5 empty] | Fax (541) 929-5277
Philomath, OR 97370 | [6 empty] | www.wetlabs.com
Scattering Meter Calibration Sheet
2004-06-29 07:48 |  | Customer: | Emmanual Boss
Wavelength: 555 |  | S/N#: | BB9-132 |  | Job #: | 404020 | Tech: | K.C.
Use the following equation to obtain "scaled" output values:
b(qc) m-1 sr-1 = Scale Factor x (Output - Dark Counts)
• | Scale Factor for 555 nm |  | = | 1.8787e-05 | (counts)
• | Output |  | = | meter output | (counts)
• | Dark Counts |  | = | 82.294 | (counts)
| Instrument Resolution |  | = | 4.9971 | (counts) | 9.3881e-05 | (m-1 sr-1)
Definitions:
• Scale Factor: Calibration scale factor, b(qc)/counts. Refer to User's Guide for derivation.
• Output: Measured signal output of the scattering meter.
• Dark Counts: Signal obtained by covering detector with black tape and submersing sensor in water.
Instrument Resolution: Standard deviation of 1 minute of collected data.
1. For each scattering centroid angle (qc), we compute the weighting function W(q,qc), by numerical integration of sample volume elements according to the sensor geometry.
2. We determine scattering phase functions, b(q,l)/b(l), for the microsphere calibration particles by weighing volume scattering functions computed from Mie theory according to the known size distribution of the microsphere polydispersion and normalizing to total scattering.
3. By convolving W(q,qc) with b(q,l)/b(l), we then compute the normalized volume scattering coefficient for each measurement angle, b(q,l)/b(l), with units of 1/sr: b(qc)/b for 2.00 micron diameter beads.
Wavelength, l (nm)
|  | Angle, qc | 555
|  | 117 | 0.005344
4. We then experimentally obtain raw scattering counts simultaneously with attenuation coefficients (Cp, using an ac-9) for a concentration series of the bead polydispersion. Absorption by the calibration particles is assumed negligible.
Results for this instrument follow:
raw counts
|  | Cp555 | 117
|  |  | 82.294 | Dark Counts: obtained by covering detector with black tape and submersing sensor in water.
| Clean water | 0.001 | 171.714
|  | 0.23 | 144.636
|  | 0.405 | 184.88
|  | 1.22 | 435.132
|  | 1.817 | 619.394
|  | 2.414 | 787.371
|  | 3.058 | 979.618
5. We obtain b/counts from the slope of a linear regression between Cp (equivalent to b for the beads) and counts:
117
| slope, intercept | 0.0035155 | -0.36232
|  | 0.00015761 | 0.0891
| r | 0.99005 | 0.12804
6. Multiplying the experimental b/counts by the theoretical b(qc)/b yields the calibration scaling factor, S:
S
| 117 | 1.8787e-05
To obtain b(qc), 1) subtract the dark counts (see step 4) from the raw counts measured, then 2) multiply by S from the table above.
7. The test also provides a measure of the inherent noise level of the instrument from the short term RMS deviation from the average number of counts. This translates into the resolution of b(qc) (minimum detectable signal change) in units of m-1 sr-1).
| Noise (counts) |  | Resolution (1/[m. ster])
| 117 | 4.9971 |  | 9.3881e-05

## 510(nm)
PO Box 518 | [5 empty] | (541) 929-5650
620 Applegate St. | [5 empty] | Fax (541) 929-5277
Philomath, OR 97370 | [6 empty] | www.wetlabs.com
Scattering Meter Calibration Sheet
2004-06-29 07:48 |  | Customer: | Emmanual Boss
Wavelength: 510 |  | S/N#: | BB9-132 |  | Job #: | 404020 | Tech: | K.C.
Use the following equation to obtain "scaled" output values:
b(qc) m-1 sr-1 = Scale Factor x (Output - Dark Counts)
• | Scale Factor for 510 nm |  | = | 8.5953e-06 | (counts)
• | Output |  | = | meter output | (counts)
• | Dark Counts |  | = | 53.941 | (counts)
| Instrument Resolution |  | = | 1.7785 | (counts) | 1.5287e-05 | (m-1 sr-1)
Definitions:
• Scale Factor: Calibration scale factor, b(qc)/counts. Refer to User's Guide for derivation.
• Output: Measured signal output of the scattering meter.
• Dark Counts: Signal obtained by covering detector with black tape and submersing sensor in water.
Instrument Resolution: Standard deviation of 1 minute of collected data.
1. For each scattering centroid angle (qc), we compute the weighting function W(q,qc), by numerical integration of sample volume elements according to the sensor geometry.
2. We determine scattering phase functions, b(q,l)/b(l), for the microsphere calibration particles by weighing volume scattering functions computed from Mie theory according to the known size distribution of the microsphere polydispersion and normalizing to total scattering.
3. By convolving W(q,qc) with b(q,l)/b(l), we then compute the normalized volume scattering coefficient for each measurement angle, b(q,l)/b(l), with units of 1/sr: b(qc)/b for 2.00 micron diameter beads.
Wavelength, l (nm)
|  | Angle, qc | 510
|  | 117 | 0.0073813
4. We then experimentally obtain raw scattering counts simultaneously with attenuation coefficients (Cp, using an ac-9) for a concentration series of the bead polydispersion. Absorption by the calibration particles is assumed negligible.
Results for this instrument follow:
raw counts
|  | Cp510 | 117
|  |  | 53.941 | Dark Counts: obtained by covering detector with black tape and submersing sensor in water.
| Clean water | 0.001 | 86.571
|  | 0.193 | 247.273
|  | 0.338 | 373.38
|  | 0.594 | 593.355
|  | 0.999 | 942.632
|  | 1.498 | 1375.212
|  | 1.988 | 1782.914
|  | 2.511 | 2243.324
5. We obtain b/counts from the slope of a linear regression between Cp (equivalent to b for the beads) and counts:
117
| slope, intercept | 0.0011645 | -0.097494
|  | 2.4307e-06 | 0.0029208
| r | 0.99997 | 0.0050087
6. Multiplying the experimental b/counts by the theoretical b(qc)/b yields the calibration scaling factor, S:
S
| 117 | 8.5953e-06
To obtain b(qc), 1) subtract the dark counts (see step 4) from the raw counts measured, then 2) multiply by S from the table above.
7. The test also provides a measure of the inherent noise level of the instrument from the short term RMS deviation from the average number of counts. This translates into the resolution of b(qc) (minimum detectable signal change) in units of m-1 sr-1).
| Noise (counts) |  | Resolution (1/[m. ster])
| 117 | 1.7785 |  | 1.5287e-05

## 532(green)
PO Box 518 | [5 empty] | (541) 929-5650
620 Applegate St. | [5 empty] | Fax (541) 929-5277
Philomath, OR 97370 | [6 empty] | www.wetlabs.com
Scattering Meter Calibration Sheet
2004-06-29 07:49 |  | Customer: | Emmanual Boss
Wavelength: 532 |  | S/N#: | BB9-132 |  | Job #: | 404020 | Tech: | K.C.
Use the following equation to obtain "scaled" output values:
b(qc) m-1 sr-1 = Scale Factor x (Output - Dark Counts)
• | Scale Factor for 532 nm |  | = | 7.8259e-06 | (counts)
• | Output |  | = | meter output | (counts)
• | Dark Counts |  | = | 77.853 | (counts)
| Instrument Resolution |  | = | 3.212 | (counts) | 2.5137e-05 | (m-1 sr-1)
Definitions:
• Scale Factor: Calibration scale factor, b(qc)/counts. Refer to User's Guide for derivation.
• Output: Measured signal output of the scattering meter.
• Dark Counts: Signal obtained by covering detector with black tape and submersing sensor in water.
Instrument Resolution: Standard deviation of 1 minute of collected data.
1. For each scattering centroid angle (qc), we compute the weighting function W(q,qc), by numerical integration of sample volume elements according to the sensor geometry.
2. We determine scattering phase functions, b(q,l)/b(l), for the microsphere calibration particles by weighing volume scattering functions computed from Mie theory according to the known size distribution of the microsphere polydispersion and normalizing to total scattering.
3. By convolving W(q,qc) with b(q,l)/b(l), we then compute the normalized volume scattering coefficient for each measurement angle, b(q,l)/b(l), with units of 1/sr: b(qc)/b for 2.00 micron diameter beads.
Wavelength, l (nm)
|  | Angle, qc | 532
|  | 117 | 0.0062564
4. We then experimentally obtain raw scattering counts simultaneously with attenuation coefficients (Cp, using an ac-9) for a concentration series of the bead polydispersion. Absorption by the calibration particles is assumed negligible.
Results for this instrument follow:
raw counts
|  | Cp532 | 117
|  |  | 77.853 | Dark Counts: obtained by covering detector with black tape and submersing sensor in water.
| Clean water | 0.001 | 250.286
|  | 0.213 | 405.879
|  | 0.373 | 549.64
|  | 0.659 | 776.935
|  | 1.114 | 1148.474
|  | 1.659 | 1588.697
|  | 2.211 | 2015.6
|  | 2.793 | 2472.265
5. We obtain b/counts from the slope of a linear regression between Cp (equivalent to b for the beads) and counts:
117
| slope, intercept | 0.0012509 | -0.31184
|  | 5.5526e-06 | 0.0076414
| r | 0.99988 | 0.011848
6. Multiplying the experimental b/counts by the theoretical b(qc)/b yields the calibration scaling factor, S:
S
| 117 | 7.8259e-06
To obtain b(qc), 1) subtract the dark counts (see step 4) from the raw counts measured, then 2) multiply by S from the table above.
7. The test also provides a measure of the inherent noise level of the instrument from the short term RMS deviation from the average number of counts. This translates into the resolution of b(qc) (minimum detectable signal change) in units of m-1 sr-1).
| Noise (counts) |  | Resolution (1/[m. ster])
| 117 | 3.212 |  | 2.5137e-05

## 660(red)
620 Applegate St. | [5 empty] | Fax (541) 929-5277
Philomath, OR 97370 | [6 empty] | www.wetlabs.com
Scattering Meter Calibration Sheet
2004-06-29 07:49 |  | Customer: | Emmanual Boss
Wavelength: 650 |  | S/N#: | BB9-132 |  | Job #: | 404020 | Tech: | K.C.
Use the following equation to obtain "scaled" output values:
b(qc) m-1 sr-1 = Scale Factor x (Output - Dark Counts)
• | Scale Factor for 660 nm |  | = | 4.0023e-06 | (counts)
• | Output |  | = | meter output | (counts)
• | Dark Counts |  | = | 43.219 | (counts)
| Instrument Resolution |  | = | 2.0597 | (counts) | 8.2435e-06 | (m-1 sr-1)
Definitions:
• Scale Factor: Calibration scale factor, b(qc)/counts. Refer to User's Guide for derivation.
• Output: Measured signal output of the scattering meter.
• Dark Counts: Signal obtained by covering detector with black tape and submersing sensor in water.
Instrument Resolution: Standard deviation of 1 minute of collected data.
1. For each scattering centroid angle (qc), we compute the weighting function W(q,qc), by numerical integration of sample volume elements according to the sensor geometry.
2. We determine scattering phase functions, b(q,l)/b(l), for the microsphere calibration particles by weighing volume scattering functions computed from Mie theory according to the known size distribution of the microsphere polydispersion and normalizing to total scattering.
3. By convolving W(q,qc) with b(q,l)/b(l), we then compute the normalized volume scattering coefficient for each measurement angle, b(q,l)/b(l), with units of 1/sr: b(qc)/b for 2.00 micron diameter beads.
Wavelength, l (nm)
|  | Angle, qc | 650
|  | 117 | 0.0031974
4. We then experimentally obtain raw scattering counts simultaneously with attenuation coefficients (Cp, using an ac-9) for a concentration series of the bead polydispersion. Absorption by the calibration particles is assumed negligible.
Results for this instrument follow:
raw counts
|  | Cp650 | 117
|  |  | 43.219 | Dark Counts: obtained by covering detector with black tape and submersing sensor in water.
| Clean water | 0 | 70.273
|  | 0.294 | 305.583
|  | 0.518 | 484.576
|  | 1.551 | 1316.667
|  | 2.325 | 1946.667
|  | 3.081 | 2537.257
|  | 3.909 | 3184.667
5. We obtain b/counts from the slope of a linear regression between Cp (equivalent to b for the beads) and counts:
117
| slope, intercept | 0.0012517 | -0.092326
|  | 3.9723e-06 | 0.0071116
| r | 0.99995 | 0.011641
6. Multiplying the experimental b/counts by the theoretical b(qc)/b yields the calibration scaling factor, S:
S
| 117 | 4.0023e-06
To obtain b(qc), 1) subtract the dark counts (see step 4) from the raw counts measured, then 2) multiply by S from the table above.
7. The test also provides a measure of the inherent noise level of the instrument from the short term RMS deviation from the average number of counts. This translates into the resolution of b(qc) (minimum detectable signal change) in units of m-1 sr-1).
| Noise (counts) |  | Resolution (1/[m. ster])
| 117 | 2.0597 |  | 8.2435e-06

## mike's data
wl (nm) | angle | diam (um) | sd (um) | np | beta/b | %diff wrt 1.96beads |  | Qb
440 | 100 | 2 | 0.08 | 1.6158 | 0.0076801 | 2.0781
440 | 125 | 2 | 0.08 | 1.6158 | 0.011368 | -0.18609 | 440 | 2.0853 |  | 0.71817
440 | 150 | 2 | 0.08 | 1.6158 | 0.0094262 | -1.005 | 450 | 1.6825 | 0.80682
470 | 100 | 2 | 0.08 | 1.6087 | 0.0076382 | -1.7615 | 470 | 1.7432 | 1.0438
470 | 125 | 2 | 0.08 | 1.6087 | 0.01007 | -3.4311 | 488 | 1.6701 |  | 0.57517
470 | 150 | 2 | 0.08 | 1.6087 | 0.0078581 | -4.6445
528 | 100 | 2 | 0.08 | 1.5988 | 0.0057283 | -5.1208 | 530 |  | 0.989
528 | 125 | 2 | 0.08 | 1.5988 | 0.0064476 | -7.1308 | 532 | 1.8239 |  | 0.62815
528 | 150 | 2 | 0.08 | 1.5988 | 0.0045044 | -8.4121
660 | 100 | 2 | 0.08 | 1.5861 | 0.0031669 | -3.2619
660 | 125 | 2 | 0.08 | 1.5861 | 0.0029804 | -5.2038
660 | 150 | 2 | 0.08 | 1.5861 | 0.0020224 | -1.3563
440 | 117 | 2 | 0.08 | 1.6158 | 0.010592 | 0.41786
470 | 117 | 2 | 0.08 | 1.6087 | 0.0097142 | -3.2813
528 | 117 | 2 | 0.08 | 1.5988 | 0.006417 | -6.0232 | 650 | 2.9036
660 | 117 | 2 | 0.08 | 1.5861 | 0.0030697 | -3.8581 | 660 | 2.9712 | 1.0233
[7 empty] | 676 | 3.0936 |  | 1.04
700 | 117 | 2 | 0.08 | 1.5837 | 0.0027063 | -3.9753 | 700 | 3.2446 | 1.1174
880 | 117 | 2 | 0.08 | 1.577 | 0.0019635 | -1.6829 | 880 | 3.6481 | 1.2564
412 | 117 | 2 | 0.08 | 1.6241 | 0.0099378
440 | 117 | 2 | 0.08 | 1.6158 | 0.010592
488 | 117 | 2 | 0.08 | 1.6052 | 0.008712
510 | 117 | 2 | 0.08 | 1.6014 | 0.0073813
532 | 117 | 2 | 0.08 | 1.5982 | 0.0062564
555 | 117 | 2 | 0.08 | 1.5953 | 0.005344
650 | 117 | 2 | 0.08 | 1.5868 | 0.0031974
676 | 117 | 2 | 0.08 | 1.5851 | 0.0029298
715 | 117 | 2 | 0.08 | 1.5829 | 0.0026085

## 676(nm)
PO Box 518 | [5 empty] | (541) 929-5650
620 Applegate St. | [5 empty] | Fax (541) 929-5277
Philomath, OR 97370 | [6 empty] | www.wetlabs.com
Scattering Meter Calibration Sheet
2026-07-28 20:17 |  | Customer: | Emmanual Boss
Wavelength:676 |  | S/N#: | BB9-132 |  | Job #: | 404020 | Tech: | K.C.
Use the following equation to obtain "scaled" output values:
b(qc) m-1 sr-1 = Scale Factor x (Output - Dark Counts)
• | Scale Factor for 676 nm |  | = | 3.6511e-06 | (counts)
• | Output |  | = | meter output | (counts)
• | Dark Counts |  | = | 52.313 | (counts)
| Instrument Resolution |  | = | 1.0586 | (counts) | 3.865e-06 | (m-1 sr-1)
Definitions:
• Scale Factor: Calibration scale factor, b(qc)/counts. Refer to User's Guide for derivation.
• Output: Measured signal output of the scattering meter.
• Dark Counts: Signal obtained by covering detector with black tape and submersing sensor in water.
Instrument Resolution: Standard deviation of 1 minute of collected data.
1. For each scattering centroid angle (qc), we compute the weighting function W(q,qc), by numerical integration of sample volume elements according to the sensor geometry.
2. We determine scattering phase functions, b(q,l)/b(l), for the microsphere calibration particles by weighing volume scattering functions computed from Mie theory according to the known size distribution of the microsphere polydispersion and normalizing to total scattering.
3. By convolving W(q,qc) with b(q,l)/b(l), we then compute the normalized volume scattering coefficient for each measurement angle, b(q,l)/b(l), with units of 1/sr: b(qc)/b for 2.00 micron diameter beads.
Wavelength, l (nm)
|  | Angle, qc | 676
|  | 117 | 0.0029298
4. We then experimentally obtain raw scattering counts simultaneously with attenuation coefficients (Cp, using an ac-9) for a concentration series of the bead polydispersion. Absorption by the calibration particles is assumed negligible.
Results for this instrument follow:
raw counts
|  | Cp676 | 117
|  |  | 52.313 | Dark Counts: obtained by covering detector with black tape and submersing sensor in water.
| Clean water | 0.002 | 77.364
|  | 0.296 | 319.833
|  | 0.525 | 503.818
|  | 1.575 | 1348.424
|  | 2.364 | 1986.833
|  | 3.139 | 2607.943
|  | 3.99 | 3272.909
5. We obtain b/counts from the slope of a linear regression between Cp (equivalent to b for the beads) and counts:
117
| slope, intercept | 0.0012462 | -0.10239
|  | 3.0775e-06 | 0.0056556
| r | 0.99997 | 0.0092421
6. Multiplying the experimental b/counts by the theoretical b(qc)/b yields the calibration scaling factor, S:
S
| 117 | 3.6511e-06
To obtain b(qc), 1) subtract the dark counts (see step 4) from the raw counts measured, then 2) multiply by S from the table above.
7. The test also provides a measure of the inherent noise level of the instrument from the short term RMS deviation from the average number of counts. This translates into the resolution of b(qc) (minimum detectable signal change) in units of m-1 sr-1).
| Noise (counts) |  | Resolution (1/[m. ster])
| 117 | 1.0586 |  | 3.865e-06

## 412(nm)
PO Box 518 | [5 empty] | (541) 929-5650
620 Applegate St. | [5 empty] | Fax (541) 929-5277
Philomath, OR 97370 | [6 empty] | www.wetlabs.com
Scattering Meter Calibration Sheet
2004-06-29 07:52 |  | Customer: | Emmanual Boss
Wavelength: 412 |  | S/N#: | BB9-132 |  | Job #: | 404020 | Tech: | K.C.
Use the following equation to obtain "scaled" output values:
b(qc) m-1 sr-1 = Scale Factor x (Output - Dark Counts)
• | Scale Factor for 412 nm |  | = | 1.1664e-05 | (counts)
• | Output |  | = | meter output | (counts)
• | Dark counts |  | = | 50.091 | (counts)
| Instrument Resolution |  | = | 2.617 | (counts) | 3.0526e-05 | (m-1 sr-1)
Definitions:
• Scale Factor: Calibration scale factor, b(qc)/counts. Refer to User's Guide for derivation.
• Output: Measured signal output of the scattering meter.
• Dark Counts: Signal obtained by covering detector with black tape and submersing sensor in water.
Instrument Resolution: Standard deviation of 1 minute of collected data.
1. For each scattering centroid angle (qc), we compute the weighting function W(q,qc), by numerical integration of sample volume elements according to the sensor geometry.
2. We determine scattering phase functions, b(q,l)/b(l), for the microsphere calibration particles by weighing volume scattering functions computed from Mie theory according to the known size distribution of the microsphere polydispersion and normalizing to total scattering
3. By convolving W(q,qc) with b(q,l)/b(l), we then compute the normalized volume scattering coefficient for each measurement angle, b(q,l)/b(l), with units of 1/sr: b(qc)/b for 2.00 micron diameter beads.
Wavelength, l (nm)
|  | Angle, qc | 412
|  | 117 | 0.0099378
4. We then experimentally obtain raw scattering counts simultaneously with attenuation coefficients (Cp, using an ac-9) for a concentration series of the bead polydispersion. Absorption by the calibration particles is assumed negligible.
Results for this instrument follow:
raw counts
|  | Cp412 | 117
|  |  | 50.091 | Dark Counts: obtained by covering detector with black tape and submersing sensor in water.
| Clean water | 0.003 | 78.71
|  | 0.171 | 212.432
|  | 0.291 | 316.941
|  | 0.5 | 497.081
|  | 0.837 | 780.788
|  | 1.239 | 1135.167
|  | 1.64 | 1474.257
|  | 2.075 | 1832.316
5. We obtain b/counts from the slope of a linear regression between Cp (equivalent to b for the beads) and counts:
117
| slope, intercept | 0.0011737 | -0.083887
|  | 4.1347e-06 | 0.00409
| r | 0.99993 | 0.006947
6. Multiplying the experimental b/counts by the theoretical b(qc)/b yields the calibration scaling factor, S:
S
| 117 | 1.1664e-05
To obtain b(qc), 1) subtract the dark counts (see step 4) from the raw counts measured, then 2) multiply by S from the table above.
7. The test also provides a measure of the inherent noise level of the instrument from the short term RMS deviation from the average number of counts. This translates into the resolution of b(qc) (minimum detectable signal change) in units of m-1 sr-1).
| Noise (counts) |  | Resolution (1/[m. ster])
| 117 | 2.617 |  | 3.0526e-05

## 715(nm)
PO Box 518 | [5 empty] | (541) 929-5650
620 Applegate St. | [5 empty] | Fax (541) 929-5277
Philomath, OR 97370 | [6 empty] | www.wetlabs.com
Scattering Meter Calibration Sheet
2004-06-29 07:50 |  | Customer: | Emmanual Boss
Wavelength: 715 |  | S/N#: | BB9-132 |  | Job #: | 404020 | Tech: | K.C.
Use the following equation to obtain "scaled" output values:
b(qc) m-1 sr-1 = Scale Factor x (Output - Dark Counts)
• | Scale Factor for 715 nm |  | = | 3.3396e-06 | (counts)
• | Output |  | = | meter output | (counts)
• | Dark Counts |  | = | 49.25 | (counts)
| Instrument Resolution |  | = | 0.9928 | (counts) | 3.3155e-06 | (m-1 sr-1)
Definitions:
• Scale Factor: Calibration scale factor, b(qc)/counts. Refer to User's Guide for derivation.
• Output: Measured signal output of the scattering meter.
• Dark Counts: Signal obtained by covering detector with black tape and submersing sensor in water.
Instrument Resolution: Standard deviation of 1 minute of collected data.
1. For each scattering centroid angle (qc), we compute the weighting function W(q,qc), by numerical integration of sample volume elements according to the sensor geometry.
2. We determine scattering phase functions, b(q,l)/b(l), for the microsphere calibration particles by weighing volume scattering functions computed from Mie theory according to the known size distribution of the microsphere polydispersion and normalizing to total scattering.
3. By convolving W(q,qc) with b(q,l)/b(l), we then compute the normalized volume scattering coefficient for each measurement angle, b(q,l)/b(l), with units of 1/sr: b(qc)/b for 2.00 micron diameter beads.
Wavelength, l (nm)
|  | Angle, qc | 715
|  | 117 | 0.0026085
4. We then experimentally obtain raw scattering counts simultaneously with attenuation coefficients (Cp, using an ac-9) for a concentration series of the bead polydispersion. Absorption by the calibration particles is assumed negligible.
Results for this instrument follow:
raw counts
|  | Cp715 | 117
|  |  | 49.25 | Dark Counts: obtained by covering detector with black tape and submersing sensor in water.
| Clean water | 0.012 | 76.909
|  | 0.323 | 315.194
|  | 0.556 | 496.818
|  | 1.634 | 1338.03
|  | 2.449 | 1983.361
|  | 3.244 | 2575.429
|  | 4.109 | 3286.485
5. We obtain b/counts from the slope of a linear regression between Cp (equivalent to b for the beads) and counts:
117
| slope, intercept | 0.0012803 | -0.081144
|  | 5.1583e-06 | 0.0094533
| r | 0.99992 | 0.01549
6. Multiplying the experimental b/counts by the theoretical b(qc)/b yields the calibration scaling factor, S:
S
| 117 | 3.3396e-06
To obtain b(qc), 1) subtract the dark counts (see step 4) from the raw counts measured, then 2) multiply by S from the table above.
7. The test also provides a measure of the inherent noise level of the instrument from the short term RMS deviation from the average number of counts. This translates into the resolution of b(qc) (minimum detectable signal change) in units of m-1 sr-1).
| Noise (counts) |  | Resolution (1/[m. ster])
| 117 | 0.9928 |  | 3.3155e-06

## Revhistory
ECO BB Calibration Sheets
Revision History
Revision | Date | Description | Originator
A | 2002-12-09 | New document (DCR 260) | D. Whiteman
A1 | 2003-02-05 | Add Chlorophyll char sheet for BB2F (DCR xxx) | K. Coe
B | 2003-03-11 | Finalize Chlorophyll char sheet (DCR 274) | K. Coe
C | 2003-05-07 | Update 1.96µm values (DCR 298) | M. Twardowski, D. Whiteman
D | 2003-06-11 | Add simplified char sheet, remove temp correction equation (DCR 306) | I. Walsh
E | 2003-06-23 | Standardize wavelengths to 450, 470, 530, 650 700, 880 (DCR 311) | D. Whiteman
F | 2003-06-30 | Separate scale factor onto separate sheet (DCR 312) | D. Whiteman
G | 2003-08-26 | Update Cal Sheet Template (DCR 330) | D. Hankins, H. Van Zee, I. Walsh
H | 2003-10-27 | Replace beta/b values per M. Twardowski (DCR 335) | D. Whiteman
I | 2004-06-02 | Change CWO to Dark count value (DCR 393) | M. Twardowski